## reprocess
year | station_code | channel id | start | duration [s] | f_min [Hz] | f_max [Hz] | windows | gain [dB] | phase [°] | normalized error | temperature [°C] | f_max [Hz] | gain | delta_t | findpeaks_calib | sensor_model | sensor_id | Kg | Rg | CDR | Rd | Km | M | Rm | f0 | lambda_x | preamp_model | preamp_id | Kp | Kc | Rp | digitizer_model | digitizer_id | Si | Kd | So | Rc | Sg | lambda0 | Sd | kappa_g | kappa_c | NCALPER | NCALIB0 | gain0 | NCALIB | gain_expected | fc | findpeaks_room | gain_room | findpeaks_deviation | findpeaks_room_deviation | gain_deviation | gain_room_deviation | findpeaks_drift | findpeaks_room_drift | gain_drift | gain_room_drift
2015 | YKAB0 | .3292..HHZ | 2015-11-23 21:00 | 218.975 | 0.2 | 20 | 86 | -10.8 | -72.5 | 1e-06 | -18.4 | 20 | 0.2884 | -18.4 | 0.003569 | S-13 | 2340 | 618.3 | 3754 | 6132 | 6132 | 0.1975 | 5 | 23 | 1 | 0.0075 | MAN-ELP-0110 | G18331 | 40 | 8.692 | 200 | CMG-DM24 | A3292 | 2.868e-06 | 2 | 4.24e-07 | 25.6558 | 383.5136 | 0.62027 | 697350.0697 | -0.0016923 | -0.0022977 | 0.25 | 0.0037194 | 0.27786 | 0.0036801 | 0.28083 | 4 | 0.0037154 | 0.27657 | 3.0191 | -0.96025 | 2.698 | -1.5159 | 0 | 0 | 0 | 0
2015 | YKAB1 | .3085..HHZ | 2015-11-23 19:26 | 287.375 | 0.2 | 20 | 113 | -10.83 | -72.3 | 1.6e-05 | -19.8 | 20 | 0.28741 | -19.8 | 0.003565 | S-13 | 2338 | 605.92 | 3729 | 5971 | 5971 | 0.1975 | 5 | 23 | 1 | 0.0212 | MAN-ELP-0110 | G18317 | 40 | 8.692 | 200 | CMG-DM24 | A3085 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 372.9844 | 0.61557 | 696864.1115 | -0.0017108 | -0.0023162 | 0.25 | 0.003827 | 0.27005 | 0.0037838 | 0.27313 | 4 | 0.0037224 | 0.27472 | 5.7828 | 1.6227 | 5.2279 | 0.58171 | 0 | 0 | 0 | 0
2015 | YKAB2 | .3087..HHZ | 2015-11-23 19:36 | 287.575 | 0.2 | 20 | 114 | -10.7 | -72.2 | 2e-05 | -19.8 | 20 | 0.29174 | -19.8 | 0.003521 | S-13 | 2342 | 613.32 | 3831 | 6158 | 6158 | 0.1975 | 5 | 23 | 1 | 0.0275 | MAN-ELP-0110 | G18338 | 40 | 8.692 | 200 | CMG-DM24 | A3087 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 378.0984 | 0.61648 | 696864.1115 | -0.0017072 | -0.0023126 | 0.25 | 0.0037753 | 0.27375 | 0.0037332 | 0.27683 | 4 | 0.0036765 | 0.27886 | 5.6833 | 1.5189 | 5.3852 | 0.73199 | 0 | 0 | 0 | 0
2015 | YKAB3 | .3080..HHZ | 2015-11-23 19:49 | 292.975 | 0.2 | 20 | 116 | -10.94 | -72.4 | 1.3e-05 | -19.8 | 20 | 0.28379 | -19.8 | 0.00365 | S-13 | 2343 | 607.6 | 3764 | 5910 | 5910 | 0.1975 | 5 | 23 | 1 | 0.0055 | MAN-ELP-0110 | G18318 | 40 | 8.692 | 200 | CMG-DM24 | A3080 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 371.1925 | 0.61092 | 696864.1115 | -0.0017291 | -0.0023344 | 0.25 | 0.0038455 | 0.26875 | 0.0037993 | 0.27202 | 4 | 0.0038112 | 0.27126 | 3.9294 | -0.3125 | 4.3288 | -0.27772 | 0 | 0 | 0 | 0
2015 | YKAB4 | .3180..HHZ | 2015-11-23 23:30 | 288.975 | 0.2 | 20 | 114 | -10.69 | -72.3 | 1.4e-05 | -18.9 | 20 | 0.29208 | -18.9 | 0.003534 | S-13 | 831 | 619.12 | 3646 | 6239 | 6239 | 0.1975 | 5 | 23 | 1 | 0.0195 | MAN-ELP-0110 | G18326 | 40 | 8.692 | 200 | CMG-DM24 | A3180 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 390.7627 | 0.63116 | 697107.0059 | -0.0016495 | -0.0022549 | 0.25 | 0.0036516 | 0.28302 | 0.0036194 | 0.28554 | 4 | 0.0036829 | 0.27977 | 2.3595 | -1.7557 | 2.2913 | -2.02 | 0 | 0 | 0 | 0
2015 | YKAB6 | .3036..HHZ | 2015-11-23 20:13 | 292.975 | 0.2 | 20 | 116 | -10.55 | -72 | 1.5e-05 | -19.1 | 20 | 0.29682 | -19.1 | 0.003461 | S-13 | 2345 | 620.79 | 3746 | 6277 | 6277 | 0.1975 | 5 | 23 | 1 | 0.0226 | MAN-ELP-0110 | G18335 | 40 | 8.692 | 200 | CMG-DM24 | A3036 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 388.7757 | 0.62626 | 697107.0059 | -0.0016688 | -0.0022741 | 0.25 | 0.0036703 | 0.28158 | 0.003635 | 0.28431 | 4 | 0.0036084 | 0.28418 | 4.7877 | 0.73229 | 4.4024 | -0.044451 | 0 | 0 | 0 | 0
2015 | YKAB7 | .3086..HHZ | 2015-11-23 20:25 | 288.975 | 0.2 | 20 | 114 | -10.04 | 107.9 | 1.6e-05 | -19.1 | 20 | 0.31477 | -19.1 | 0.003269 | S-13 | 2347 | 639.15 | 3957 | 6665 | 6665 | 0.1975 | 5 | 23 | 1 | 0.0242 | MAN-ELP-0110 | G18337 | 40 | 8.692 | 200 | CMG-DM24 | A3086 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 401.0483 | 0.62747 | 697107.0059 | -0.001664 | -0.0022694 | 0.25 | 0.003558 | 0.29047 | 0.0035245 | 0.29323 | 4 | 0.0034082 | 0.30137 | 7.2489 | 3.2983 | 7.3489 | 2.7766 | 0 | 0 | 0 | 0
2015 | YKAB8 | .3028..HHZ | 2015-11-23 20:36 | 287.975 | 0.2 | 20 | 114 | -10.6 | -71.9 | 1.4e-05 | -19.1 | 20 | 0.29512 | -19.1 | 0.003481 | S-13 | 2493 | 621.46 | 3761 | 6296 | 6296 | 0.1975 | 5 | 23 | 1 | 0.0234 | MAN-ELP-0110 | G18336 | 40 | 8.692 | 200 | CMG-DM24 | A4549 | 2.873e-06 | 2 | 4.24e-07 | 25.6558 | 389.0536 | 0.62603 | 696136.4427 | -0.0016697 | -0.002275 | 0.25 | 0.0036728 | 0.28139 | 0.0036374 | 0.28413 | 4 | 0.0036293 | 0.28255 | 4.2991 | 0.22286 | 3.8699 | -0.55423 | 0 | 0 | 0 | 0
2015 | YKAB9 | .3072..HHZ | 2015-11-23 20:47 | 287.375 | 0.2 | 20 | 113 | -10.93 | -72.3 | 1.4e-05 | -19.1 | 20 | 0.28412 | -19.1 | 0.003611 | S-13 | 2348 | 601.74 | 3754 | 5886 | 5886 | 0.1975 | 5 | 23 | 1 | 0.0207 | MAN-ELP-0110 | G18320 | 40 | 8.692 | 200 | CMG-DM24 | A3072 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.411 | 0.61058 | 698080.2792 | -0.0017304 | -0.0023358 | 0.25 | 0.0038783 | 0.26647 | 0.0038315 | 0.26973 | 4 | 0.0037648 | 0.27202 | 5.7551 | 1.7409 | 5.3347 | 0.84817 | 0 | 0 | 0 | 0
2015 | YKAR1 | .3079..HHZ | 2015-11-23 16:28 | 287.975 | 0.2 | 20 | 114 | -10.37 | -72.4 | 9.5e-06 | -23.9 | 20 | 0.30304 | -23.9 | 0.003418 | S-13 | 1192 | 619.2 | 3774 | 6249 | 6249 | 0.1975 | 5 | 23 | 1 | 0.0233 | MAN-ELP-0110 | G18321 | 40 | 8.692 | 200 | CMG-DM24 | A3079 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0502 | 0.62347 | 697593.3031 | -0.0016798 | -0.0022851 | 0.25 | 0.0036936 | 0.2798 | 0.0036565 | 0.28264 | 4 | 0.0036002 | 0.28689 | 6.523 | 1.541 | 7.2184 | 1.504 | 0 | 0 | 0 | 0
2015 | YKAR2 | .3076..HHZ | 2015-11-23 17:05 | 287.675 | 0.2 | 20 | 114 | -10.32 | -72 | 8.9e-06 | -22.7 | 20 | 0.30479 | -22.7 | 0.003375 | S-13 | 832 | 628.26 | 3591 | 6355 | 6355 | 0.1975 | 5 | 23 | 1 | 0.0112 | MAN-ELP-0110 | G18332 | 40 | 8.692 | 200 | CMG-DM24 | A3076 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 401.4269 | 0.63895 | 697593.3031 | -0.0016189 | -0.0022243 | 0.25 | 0.0035522 | 0.29094 | 0.0035252 | 0.29316 | 4 | 0.0035458 | 0.28936 | 4.2618 | -0.58452 | 3.9658 | -1.297 | 0 | 0 | 0 | 0
2015 | YKAR3 | .3073..HHZ | 2015-11-23 17:29 | 154 | 0.2 | 20 | 60 | -11.63 | -72.6 | 1.1e-06 | -22.7 | 20 | 0.26212 | -22.7 | 0.003827 | S-13 | 822 | 619.92 | 3790 | 6255 | 6255 | 0.1975 | 5 | 23 | 1 | 0.0219 | MAN-ELP-0110 | G18334 | 40 | 8.692 | 200 | CMG-DM24 | A3073 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0229 | 0.6227 | 697593.3031 | -0.0016828 | -0.0022881 | 0.25 | 0.0036939 | 0.27978 | 0.0036563 | 0.28265 | 4 | 0.0040207 | 0.24885 | -4.6681 | -9.9665 | -7.2645 | -11.9589 | 0 | 0 | 0 | 0
2015 | YKAR4 | .3081..HHZ | 2015-11-23 17:57 | 167.675 | 0.2 | 20 | 66 | -10.43 | 107.6 | 1.9e-05 | -22.7 | 20 | 0.30095 | -22.7 | 0.00341 | S-13 | 2474 | 626.6 | 3619 | 6407 | 6407 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18333 | 40 | 8.692 | 200 | CMG-DM24 | A3081 | 2.864e-06 | 2 | 4.24e-07 | 25.6558 | 400.4215 | 0.63904 | 698324.0223 | -0.0016186 | -0.0022239 | 0.25 | 0.0035573 | 0.29052 | 0.0035304 | 0.29273 | 4 | 0.0035826 | 0.28572 | 3.4116 | -1.4778 | 2.809 | -2.3953 | 0 | 0 | 0 | 0
2015 | YKAR5 | .3077..HHZ | 2015-11-23 18:16 | 288.375 | 0.2 | 20 | 114 | -10.81 | 107.6 | 1.6e-05 | -21 | 20 | 0.28807 | -21 | 0.003557 | S-13 | 2349 | 615.83 | 3762 | 6164 | 6164 | 0.1975 | 5 | 23 | 1 | 0.0205 | MAN-ELP-0110 | G18325 | 40 | 8.692 | 200 | CMG-DM24 | A3077 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 382.4276 | 0.621 | 697107.0059 | -0.0016895 | -0.0022948 | 0.25 | 0.0037312 | 0.27698 | 0.0036923 | 0.2799 | 4 | 0.0037236 | 0.27458 | 3.6635 | -0.84795 | 2.9188 | -1.9009 | 0 | 0 | 0 | 0
2015 | YKAR6 | .3038..HHZ | 2015-11-23 18:36 | 292.975 | 0.2 | 20 | 116 | -10.5 | -72.3 | 1.4e-05 | -21 | 20 | 0.29854 | -21 | 0.003388 | S-13 | 2350 | 616.95 | 3661 | 6211 | 6211 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18324 | 40 | 8.692 | 200 | CMG-DM24 | A3038 | 2.875e-06 | 2 | 4.24e-07 | 25.6558 | 388.156 | 0.62915 | 695652.1739 | -0.0016574 | -0.0022628 | 0.25 | 0.0036838 | 0.28054 | 0.0036501 | 0.28313 | 4 | 0.0035467 | 0.28456 | 7.1819 | 2.8352 | 5.4417 | 0.50389 | 0 | 0 | 0 | 0
2015 | YKAR7 | .3074..HHZ | 2015-11-23 18:52 | 292.375 | 0.2 | 20 | 115 | -10.37 | -71.9 | 1e-05 | -21 | 20 | 0.30304 | -21 | 0.003397 | S-13 | 2911 | 618.86 | 3577 | 6252 | 6252 | 0.1975 | 5 | 23 | 1 | 0.0248 | MAN-ELP-0110 | G18329 | 40 | 8.692 | 200 | CMG-DM24 | A3074 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 393.6426 | 0.63608 | 697836.7062 | -0.0016302 | -0.0022356 | 0.25 | 0.0036211 | 0.2854 | 0.003592 | 0.28771 | 4 | 0.0035561 | 0.28885 | 5.4293 | 1.0005 | 5.3272 | 0.39476 | 0 | 0 | 0 | 0
2015 | YKAR8 | .2165..HHZ | 2015-11-23 19:03 | 292.375 | 0.2 | 20 | 115 | -11.29 | -72.6 | 1.3e-05 | -19.8 | 20 | 0.27258 | -19.8 | 0.003788 | S-13 | 1466 | 600.25 | 3720 | 5860 | 5860 | 0.1975 | 5 | 23 | 1 | 0.0211 | MAN-ELP-0110 | G18330 | 40 | 8.692 | 200 | CMG-DM24 | A2165 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.1675 | 0.61169 | 698080.2792 | -0.0017261 | -0.0023314 | 0.25 | 0.0038809 | 0.2663 | 0.0038347 | 0.2695 | 4 | 0.0039553 | 0.26055 | 1.2183 | -3.1433 | 1.1429 | -3.323 | 0 | 0 | 0 | 0
2015 | YKAR9 | .3291..HHZ | 2015-11-23 19:14 | 292.975 | 0.2 | 20 | 116 | -10.7 | -72.1 | 1.6e-05 | -19.8 | 20 | 0.29174 | -19.8 | 0.003515 | S-13 | 954 | 609.74 | 3688 | 6046 | 6046 | 0.1975 | 5 | 23 | 1 | 0.021 | MAN-ELP-0110 | G18322 | 40 | 8.692 | 200 | CMG-DM24 | A3291 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 378.7228 | 0.62112 | 697836.7062 | -0.001689 | -0.0022943 | 0.25 | 0.0037638 | 0.27458 | 0.0037246 | 0.27747 | 4 | 0.0036702 | 0.27886 | 5.6266 | 1.4596 | 5.1423 | 0.49984 | 0 | 0 | 0 | 0
2016 | YKAB0 | .3292..HHZ | 2016-11-14 20:54 | 286.975 | 0.2 | 20 | 113 | -11.1 | -73.7 | 1.5e-05 | -4 | 20 | 0.27861 | -4 | 0.0037163 | S-13 | 2340 | 618.3 | 3754 | 6132 | 6132 | 0.1975 | 5 | 23 | 1 | 0.0075 | MAN-ELP-0110 | G18331 | 40 | 8.692 | 200 | CMG-DM24 | A3292 | 2.868e-06 | 2 | 4.24e-07 | 25.6558 | 383.5136 | 0.62027 | 697350.0697 | -0.0016923 | -0.0022977 | 0.25 | 0.0037194 | 0.27786 | 0.0036801 | 0.28083 | 4 | 0.0037494 | 0.27613 | -0.98353 | -1.8843 | -0.78848 | -1.6734 | -4.0026 | -0.92405 | -3.4865 | -0.15757
2016 | YKAB1 | .3085..HHZ | 2016-11-14 19:39 | 292.975 | 0.2 | 20 | 116 | -11.06 | 107.3 | 1.6e-05 | -4.4 | 20 | 0.2799 | -4.4 | 0.0036737 | S-13 | 2338 | 605.92 | 3729 | 5971 | 5971 | 0.1975 | 5 | 23 | 1 | 0.0212 | MAN-ELP-0110 | G18317 | 40 | 8.692 | 200 | CMG-DM24 | A3085 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 372.9844 | 0.61557 | 696864.1115 | -0.0017108 | -0.0023162 | 0.25 | 0.003827 | 0.27005 | 0.0037838 | 0.27313 | 4 | 0.0037097 | 0.27715 | 2.91 | 1.9573 | 2.4781 | 1.4726 | -2.8728 | 0.33465 | -2.7498 | 0.89088
2016 | YKAB2 | .3087..HHZ | 2016-11-14 19:46 | 293.375 | 0.2 | 20 | 116 | -10.95 | 107.4 | 1.9e-05 | -4.4 | 20 | 0.28347 | -4.4 | 0.0036528 | S-13 | 2342 | 613.32 | 3831 | 6158 | 6158 | 0.1975 | 5 | 23 | 1 | 0.0275 | MAN-ELP-0110 | G18338 | 40 | 8.692 | 200 | CMG-DM24 | A3087 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 378.0984 | 0.61648 | 696864.1115 | -0.0017072 | -0.0023126 | 0.25 | 0.0037753 | 0.27375 | 0.0037332 | 0.27683 | 4 | 0.0036886 | 0.28068 | 2.1528 | 1.1927 | 2.3952 | 1.3905 | -3.5305 | -0.32613 | -2.99 | 0.65848
2016 | YKAB3 | .3080..HHZ | 2016-11-14 20:00 | 159.975 | 0.2 | 20 | 63 | -11.07 | 107.4 | 2.6e-06 | -4 | 20 | 0.27958 | -4 | 0.0036803 | S-13 | 2343 | 607.6 | 3764 | 5910 | 5910 | 0.1975 | 5 | 23 | 1 | 0.0055 | MAN-ELP-0110 | G18318 | 40 | 8.692 | 200 | CMG-DM24 | A3080 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 371.1925 | 0.61092 | 696864.1115 | -0.0017291 | -0.0023344 | 0.25 | 0.0038455 | 0.26875 | 0.0037993 | 0.27202 | 4 | 0.0037131 | 0.27708 | 3.1319 | 2.2678 | 2.779 | 1.8622 | -0.79752 | 2.5803 | -1.5498 | 2.1399
2016 | YKAB4 | .3180..HHZ | 2016-11-14 20:07 | 288.475 | 0.2 | 20 | 114 | -10.78 | 107.5 | 1.2e-05 | -4 | 20 | 0.28907 | -4 | 0.0035802 | S-13 | 831 | 619.12 | 3646 | 6239 | 6239 | 0.1975 | 5 | 23 | 1 | 0.0195 | MAN-ELP-0110 | G18326 | 40 | 8.692 | 200 | CMG-DM24 | A3180 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 390.7627 | 0.63116 | 697107.0059 | -0.0016495 | -0.0022549 | 0.25 | 0.0036516 | 0.28302 | 0.0036194 | 0.28554 | 4 | 0.0036121 | 0.28649 | 1.0831 | 0.20072 | 1.2368 | 0.33379 | -1.2765 | 1.9565 | -1.0544 | 2.3538
2016 | YKAB6 | .3036..HHZ | 2016-11-14 20:16 | 288.975 | 0.2 | 20 | 114 | -10.77 | -72.4 | 1.4e-05 | -4 | 20 | 0.2894 | -4 | 0.0035826 | S-13 | 2345 | 620.79 | 3746 | 6277 | 6277 | 0.1975 | 5 | 23 | 1 | 0.0226 | MAN-ELP-0110 | G18335 | 40 | 8.692 | 200 | CMG-DM24 | A3036 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 388.7757 | 0.62626 | 697107.0059 | -0.0016688 | -0.0022741 | 0.25 | 0.0036703 | 0.28158 | 0.003635 | 0.28431 | 4 | 0.0036146 | 0.28682 | 1.4424 | 0.56331 | 1.7912 | 0.88324 | -3.3452 | -0.16898 | -2.6111 | 0.92769
2016 | YKAB7 | .3086..HHZ | 2016-11-14 20:24 | 288.975 | 0.2 | 20 | 114 | -10.27 | 107.5 | 1.5e-05 | -4 | 20 | 0.30655 | -4 | 0.0033768 | S-13 | 2347 | 639.15 | 3957 | 6665 | 6665 | 0.1975 | 5 | 23 | 1 | 0.0242 | MAN-ELP-0110 | G18337 | 40 | 8.692 | 200 | CMG-DM24 | A3086 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 401.0483 | 0.62747 | 697107.0059 | -0.001664 | -0.0022694 | 0.25 | 0.003558 | 0.29047 | 0.0035245 | 0.29323 | 4 | 0.0034069 | 0.30381 | 4.1903 | 3.3356 | 4.5436 | 3.6111 | -3.0586 | 0.037323 | -2.8053 | 0.83452
2016 | YKAB8 | .4549..HHZ | 2016-11-14 20:31 | 287.775 | 0.2 | 20 | 114 | -10.84 | -72.3 | 1.5e-05 | -4 | 20 | 0.28708 | -4 | 0.0035806 | S-13 | 2493 | 621.46 | 3761 | 6296 | 6296 | 0.1975 | 5 | 23 | 1 | 0.0234 | MAN-ELP-0110 | G18336 | 40 | 8.692 | 200 | CMG-DM24 | A4549 | 2.873e-06 | 2 | 4.24e-07 | 25.6558 | 389.0536 | 0.62603 | 696136.4427 | -0.0016697 | -0.002275 | 0.25 | 0.0036728 | 0.28139 | 0.0036374 | 0.28413 | 4 | 0.0036125 | 0.28452 | 1.5608 | 0.68274 | 1.0392 | 0.13788 | -2.7382 | 0.45987 | -2.8307 | 0.69212
2016 | YKAB9 | .3072..HHZ | 2016-11-14 20:45 | 288.525 | 0.2 | 20 | 114 | -11.12 | 107.3 | 1.3e-05 | -4 | 20 | 0.27797 | -4 | 0.0036907 | S-13 | 2348 | 601.74 | 3754 | 5886 | 5886 | 0.1975 | 5 | 23 | 1 | 0.0207 | MAN-ELP-0110 | G18320 | 40 | 8.692 | 200 | CMG-DM24 | A3072 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.411 | 0.61058 | 698080.2792 | -0.0017304 | -0.0023358 | 0.25 | 0.0038783 | 0.26647 | 0.0038315 | 0.26973 | 4 | 0.0037236 | 0.27549 | 3.6749 | 2.8157 | 3.0556 | 2.1363 | -2.0801 | 1.0748 | -2.2791 | 1.2881
2016 | YKAR1 | .3079..HHZ | 2016-11-14 16:51 | 288.375 | 0.2 | 20 | 114 | -10.56 | -72.7 | 1.1e-05 | -4.5 | 20 | 0.29648 | -4.5 | 0.0034643 | S-13 | 1192 | 619.2 | 3774 | 6249 | 6249 | 0.1975 | 5 | 23 | 1 | 0.0233 | MAN-ELP-0110 | G18321 | 40 | 8.692 | 200 | CMG-DM24 | A3079 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0502 | 0.62347 | 697593.3031 | -0.0016798 | -0.0022851 | 0.25 | 0.0036936 | 0.2798 | 0.0036565 | 0.28264 | 4 | 0.0034991 | 0.29351 | 5.2568 | 4.3061 | 4.8985 | 3.8459 | -1.2662 | 2.7651 | -2.3199 | 2.3419
2016 | YKAR2 | .3076..HHZ | 2016-11-14 18:07 | 293.975 | 0.2 | 20 | 116 | -10.51 | -72.3 | 1.2e-05 | -4.5 | 20 | 0.29819 | -4.5 | 0.0034509 | S-13 | 832 | 628.26 | 3591 | 6355 | 6355 | 0.1975 | 5 | 23 | 1 | 0.0112 | MAN-ELP-0110 | G18332 | 40 | 8.692 | 200 | CMG-DM24 | A3076 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 401.4269 | 0.63895 | 697593.3031 | -0.0016189 | -0.0022243 | 0.25 | 0.0035522 | 0.29094 | 0.0035252 | 0.29316 | 4 | 0.0034855 | 0.2952 | 2.1088 | 1.1265 | 1.7163 | 0.69559 | -2.153 | 1.711 | -2.2495 | 1.9926
2016 | YKAR3 | .3073..HHZ | 2016-11-14 18:12 | 291.975 | 0.2 | 20 | 115 | -11.81 | -72.9 | 1e-05 | -4.5 | 20 | 0.25674 | -4.5 | 0.0039497 | S-13 | 822 | 619.92 | 3790 | 6255 | 6255 | 0.1975 | 5 | 23 | 1 | 0.0219 | MAN-ELP-0110 | G18334 | 40 | 8.692 | 200 | CMG-DM24 | A3073 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0229 | 0.6227 | 697593.3031 | -0.0016828 | -0.0022881 | 0.25 | 0.0036939 | 0.27978 | 0.0036563 | 0.28265 | 4 | 0.0039893 | 0.25417 | -8.0239 | -9.108 | -9.1665 | -10.078 | -3.3558 | 0.85855 | -1.902 | 1.8808
2016 | YKAR4 | .3081..HHZ | 2016-11-14 18:20 | 292.375 | 0.2 | 20 | 115 | -10.64 | -72.8 | 1.3e-05 | -4.5 | 20 | 0.29376 | -4.5 | 0.0035156 | S-13 | 2474 | 626.6 | 3619 | 6407 | 6407 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18333 | 40 | 8.692 | 200 | CMG-DM24 | A3081 | 2.864e-06 | 2 | 4.24e-07 | 25.6558 | 400.4215 | 0.63904 | 698324.0223 | -0.0016186 | -0.0022239 | 0.25 | 0.0035573 | 0.29052 | 0.0035304 | 0.29273 | 4 | 0.0035509 | 0.29082 | 0.42044 | -0.57884 | 0.35319 | -0.65386 | -2.9911 | 0.899 | -2.4558 | 1.7414
2016 | YKAR5 | .3077..HHZ | 2016-11-14 18:37 | 293.375 | 0.2 | 20 | 116 | -11.01 | 107.2 | 1.8e-05 | -4.5 | 20 | 0.28151 | -4.5 | 0.003651 | S-13 | 2349 | 615.83 | 3762 | 6164 | 6164 | 0.1975 | 5 | 23 | 1 | 0.0205 | MAN-ELP-0110 | G18325 | 40 | 8.692 | 200 | CMG-DM24 | A3077 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 382.4276 | 0.621 | 697107.0059 | -0.0016895 | -0.0022948 | 0.25 | 0.0037312 | 0.27698 | 0.0036923 | 0.2799 | 4 | 0.0036876 | 0.27869 | 1.1176 | 0.12534 | 0.57611 | -0.43317 | -2.5459 | 0.97329 | -2.3427 | 1.4677
2016 | YKAR6 | .3038..HHZ | 2016-11-14 18:51 | 287.975 | 0.2 | 20 | 114 | -10.71 | -72.5 | 1.5e-05 | -4.5 | 20 | 0.29141 | -4.5 | 0.0035116 | S-13 | 2350 | 616.95 | 3661 | 6211 | 6211 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18324 | 40 | 8.692 | 200 | CMG-DM24 | A3038 | 2.875e-06 | 2 | 4.24e-07 | 25.6558 | 388.156 | 0.62915 | 695652.1739 | -0.0016574 | -0.0022628 | 0.25 | 0.0036838 | 0.28054 | 0.0036501 | 0.28313 | 4 | 0.0035468 | 0.28848 | 3.7957 | 2.8303 | 2.923 | 1.8902 | -3.3862 | -0.0049002 | -2.5187 | 1.3863
2016 | YKAR7 | .3074..HHZ | 2016-11-14 19:09 | 292.975 | 0.2 | 20 | 116 | -10.58 | -72.5 | 1.1e-05 | -4.4 | 20 | 0.2958 | -4.4 | 0.0034639 | S-13 | 2911 | 618.86 | 3577 | 6252 | 6252 | 0.1975 | 5 | 23 | 1 | 0.0248 | MAN-ELP-0110 | G18329 | 40 | 8.692 | 200 | CMG-DM24 | A3074 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 393.6426 | 0.63608 | 697836.7062 | -0.0016302 | -0.0022356 | 0.25 | 0.0036211 | 0.2854 | 0.003592 | 0.28771 | 4 | 0.0034979 | 0.2929 | 3.5668 | 2.6206 | 2.8113 | 1.8025 | -1.8625 | 1.6201 | -2.516 | 1.4077
2016 | YKAR8 | .2165..HHZ | 2016-11-14 19:21 | 292.975 | 0.2 | 20 | 116 | -11.41 | 107.1 | 1.4e-05 | -4.4 | 20 | 0.26884 | -4.4 | 0.0038004 | S-13 | 1466 | 600.25 | 3720 | 5860 | 5860 | 0.1975 | 5 | 23 | 1 | 0.0211 | MAN-ELP-0110 | G18330 | 40 | 8.692 | 200 | CMG-DM24 | A2165 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.1675 | 0.61169 | 698080.2792 | -0.0017261 | -0.0023314 | 0.25 | 0.0038809 | 0.2663 | 0.0038347 | 0.2695 | 4 | 0.0038377 | 0.26621 | 0.89491 | -0.077514 | -0.24487 | -1.2237 | -0.32336 | 3.0658 | -1.3877 | 2.0993
2016 | YKAR9 | .3291..HHZ | 2016-11-14 19:29 | 293.975 | 0.2 | 20 | 116 | -10.91 | -72.5 | 1.6e-05 | -4.4 | 20 | 0.28477 | -4.4 | 0.0036026 | S-13 | 954 | 609.74 | 3688 | 6046 | 6046 | 0.1975 | 5 | 23 | 1 | 0.021 | MAN-ELP-0110 | G18322 | 40 | 8.692 | 200 | CMG-DM24 | A3291 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 378.7228 | 0.62112 | 697836.7062 | -0.001689 | -0.0022943 | 0.25 | 0.0037638 | 0.27458 | 0.0037246 | 0.27747 | 4 | 0.0036379 | 0.28198 | 3.2746 | 2.3255 | 2.6307 | 1.6237 | -2.352 | 0.86594 | -2.5116 | 1.1239
2017 | YKAB0 | .3292..HHZ | 2017-10-10 21:34 | 292.975 | 0.2 | 20 | 116 | -11.13 | -73.3 | 1.5e-05 | 0 | 20 | 0.27765 | 0 | 0.0037525 | S-13 | 2340 | 618.3 | 3754 | 6132 | 6132 | 0.1975 | 5 | 23 | 1 | 0.0075 | MAN-ELP-0110 | G18331 | 40 | 8.692 | 200 | CMG-DM24 | A3292 | 2.868e-06 | 2 | 4.24e-07 | 25.6558 | 383.5136 | 0.62027 | 697350.0697 | -0.0016923 | -0.0022977 | 0.25 | 0.0037194 | 0.27786 | 0.0036801 | 0.28083 | 4 | 0.0037525 | 0.27765 | -1.9672 | -1.9672 | -1.1305 | -1.1305 | -4.9863 | -1.0069 | -3.8286 | 0.38532
2017 | YKAB1 | .3085..HHZ | 2017-10-10 21:19 | 292.425 | 0.2 | 20 | 115 | -11.15 | 107.2 | 1.7e-05 | 0 | 20 | 0.27701 | 0 | 0.0036899 | S-13 | 2338 | 605.92 | 3729 | 5971 | 5971 | 0.1975 | 5 | 23 | 1 | 0.0212 | MAN-ELP-0110 | G18317 | 40 | 8.692 | 200 | CMG-DM24 | A3085 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 372.9844 | 0.61557 | 696864.1115 | -0.0017108 | -0.0023162 | 0.25 | 0.003827 | 0.27005 | 0.0037838 | 0.27313 | 4 | 0.0036899 | 0.27701 | 2.4819 | 2.4819 | 1.4217 | 1.4217 | -3.3009 | 0.85916 | -3.8062 | 0.84004
2017 | YKAB2 | .3087..HHZ | 2017-10-10 21:22 | 293.525 | 0.2 | 20 | 116 | -11.04 | -72.7 | 1.9e-05 | 0 | 20 | 0.28054 | 0 | 0.0037008 | S-13 | 2342 | 613.32 | 3831 | 6158 | 6158 | 0.1975 | 5 | 23 | 1 | 0.0275 | MAN-ELP-0110 | G18338 | 40 | 8.692 | 200 | CMG-DM24 | A3087 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 378.0984 | 0.61648 | 696864.1115 | -0.0017072 | -0.0023126 | 0.25 | 0.0037753 | 0.27375 | 0.0037332 | 0.27683 | 4 | 0.0037008 | 0.28054 | 0.86705 | 0.86705 | 1.3397 | 1.3397 | -4.8163 | -0.65183 | -4.0455 | 0.60768
2017 | YKAB3 | .3080..HHZ | 2017-10-10 21:24 | 291.975 | 0.2 | 20 | 115 | -11.15 | -72.8 | 1.5e-05 | 0 | 20 | 0.27701 | 0 | 0.0037007 | S-13 | 2343 | 607.6 | 3764 | 5910 | 5910 | 0.1975 | 5 | 23 | 1 | 0.0055 | MAN-ELP-0110 | G18318 | 40 | 8.692 | 200 | CMG-DM24 | A3080 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 371.1925 | 0.61092 | 696864.1115 | -0.0017291 | -0.0023344 | 0.25 | 0.0038455 | 0.26875 | 0.0037993 | 0.27202 | 4 | 0.0037007 | 0.27701 | 2.5949 | 2.5949 | 1.8367 | 1.8367 | -1.3345 | 2.9074 | -2.4921 | 2.1144
2017 | YKAB4 | .3180..HHZ | 2017-10-10 20:45 | 292.525 | 0.2 | 20 | 116 | -10.84 | -72.6 | 1.4e-05 | -0.2 | 20 | 0.28708 | -0.2 | 0.0035677 | S-13 | 831 | 619.12 | 3646 | 6239 | 6239 | 0.1975 | 5 | 23 | 1 | 0.0195 | MAN-ELP-0110 | G18326 | 40 | 8.692 | 200 | CMG-DM24 | A3180 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 390.7627 | 0.63116 | 697107.0059 | -0.0016495 | -0.0022549 | 0.25 | 0.0036516 | 0.28302 | 0.0036194 | 0.28554 | 4 | 0.0035693 | 0.28695 | 1.4284 | 1.3845 | 0.53991 | 0.49507 | -0.93109 | 3.1402 | -1.7513 | 2.5151
2017 | YKAB6 | .3036..HHZ | 2017-10-10 21:26 | 292.425 | 0.2 | 20 | 115 | -10.87 | 107.4 | 1.5e-05 | 0 | 20 | 0.28609 | 0 | 0.0036157 | S-13 | 2345 | 620.79 | 3746 | 6277 | 6277 | 0.1975 | 5 | 23 | 1 | 0.0226 | MAN-ELP-0110 | G18335 | 40 | 8.692 | 200 | CMG-DM24 | A3036 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 388.7757 | 0.62626 | 697107.0059 | -0.0016688 | -0.0022741 | 0.25 | 0.0036703 | 0.28158 | 0.003635 | 0.28431 | 4 | 0.0036157 | 0.28609 | 0.53186 | 0.53186 | 0.62602 | 0.62602 | -4.2558 | -0.20043 | -3.7763 | 0.67047
2017 | YKAB7 | .3086..HHZ | 2017-10-10 21:27 | 291.975 | 0.2 | 20 | 115 | -10.37 | -72.6 | 1.6e-05 | 0 | 20 | 0.30304 | 0 | 0.0034207 | S-13 | 2347 | 639.15 | 3957 | 6665 | 6665 | 0.1975 | 5 | 23 | 1 | 0.0242 | MAN-ELP-0110 | G18337 | 40 | 8.692 | 200 | CMG-DM24 | A3086 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 401.0483 | 0.62747 | 697107.0059 | -0.001664 | -0.0022694 | 0.25 | 0.003558 | 0.29047 | 0.0035245 | 0.29323 | 4 | 0.0034207 | 0.30304 | 2.9447 | 2.9447 | 3.3469 | 3.3469 | -4.3042 | -0.35363 | -4.002 | 0.57035
2017 | YKAB8 | .4549..HHZ | 2017-10-10 21:30 | 292.975 | 0.2 | 20 | 116 | -10.94 | -72.5 | 1.6e-05 | 0 | 20 | 0.28379 | 0 | 0.0036488 | S-13 | 2493 | 621.46 | 3761 | 6296 | 6296 | 0.1975 | 5 | 23 | 1 | 0.0234 | MAN-ELP-0110 | G18336 | 40 | 8.692 | 200 | CMG-DM24 | A4549 | 2.873e-06 | 2 | 4.24e-07 | 25.6558 | 389.0536 | 0.62603 | 696136.4427 | -0.0016697 | -0.002275 | 0.25 | 0.0036728 | 0.28139 | 0.0036374 | 0.28413 | 4 | 0.0036488 | 0.28379 | -0.31417 | -0.31417 | -0.11743 | -0.11743 | -4.6132 | -0.53703 | -3.9873 | 0.4368
2017 | YKAB9 | .3072..HHZ | 2017-10-10 21:31 | 292.375 | 0.2 | 20 | 115 | -11.21 | 107.2 | 1.5e-05 | 0 | 20 | 0.27511 | 0 | 0.0037646 | S-13 | 2348 | 601.74 | 3754 | 5886 | 5886 | 0.1975 | 5 | 23 | 1 | 0.0207 | MAN-ELP-0110 | G18320 | 40 | 8.692 | 200 | CMG-DM24 | A3072 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.411 | 0.61058 | 698080.2792 | -0.0017304 | -0.0023358 | 0.25 | 0.0038783 | 0.26647 | 0.0038315 | 0.26973 | 4 | 0.0037646 | 0.27511 | 1.7462 | 1.7462 | 1.9932 | 1.9932 | -4.0089 | 0.0053066 | -3.3414 | 1.1451
2017 | YKAR1 | .3079..HHZ | 2017-10-10 20:42 | 291.975 | 0.2 | 20 | 115 | -10.66 | -72.8 | 1.2e-05 | -0.2 | 20 | 0.29309 | -0.2 | 0.0034884 | S-13 | 1192 | 619.2 | 3774 | 6249 | 6249 | 0.1975 | 5 | 23 | 1 | 0.0233 | MAN-ELP-0110 | G18321 | 40 | 8.692 | 200 | CMG-DM24 | A3079 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0502 | 0.62347 | 697593.3031 | -0.0016798 | -0.0022851 | 0.25 | 0.0036936 | 0.2798 | 0.0036565 | 0.28264 | 4 | 0.00349 | 0.29296 | 4.5977 | 4.5552 | 3.6978 | 3.6515 | -1.9253 | 3.0142 | -3.5207 | 2.1475
2017 | YKAR2 | .3076..HHZ | 2017-10-10 21:03 | 292.975 | 0.2 | 20 | 116 | -10.61 | -72.5 | 1.6e-05 | 0 | 20 | 0.29478 | 0 | 0.0035063 | S-13 | 832 | 628.26 | 3591 | 6355 | 6355 | 0.1975 | 5 | 23 | 1 | 0.0112 | MAN-ELP-0110 | G18332 | 40 | 8.692 | 200 | CMG-DM24 | A3076 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 401.4269 | 0.63895 | 697593.3031 | -0.0016189 | -0.0022243 | 0.25 | 0.0035522 | 0.29094 | 0.0035252 | 0.29316 | 4 | 0.0035063 | 0.29478 | 0.53727 | 0.53727 | 0.55198 | 0.55198 | -3.7246 | 1.1218 | -3.4139 | 1.849
2017 | YKAR3 | .3073..HHZ | 2017-10-10 21:05 | 293.375 | 0.2 | 20 | 116 | -11.88 | -72.9 | 1.3e-05 | 0 | 20 | 0.25468 | 0 | 0.0040484 | S-13 | 822 | 619.92 | 3790 | 6255 | 6255 | 0.1975 | 5 | 23 | 1 | 0.0219 | MAN-ELP-0110 | G18334 | 40 | 8.692 | 200 | CMG-DM24 | A3073 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0229 | 0.6227 | 697593.3031 | -0.0016828 | -0.0022881 | 0.25 | 0.0036939 | 0.27978 | 0.0036563 | 0.28265 | 4 | 0.0040484 | 0.25468 | -10.7234 | -10.7234 | -9.8956 | -9.8956 | -6.0553 | -0.75687 | -2.6311 | 2.0633
2017 | YKAR4 | .3081..HHZ | 2017-10-10 21:10 | 292.975 | 0.2 | 20 | 116 | -10.73 | -72.9 | 1.4e-05 | 0 | 20 | 0.29074 | 0 | 0.0035433 | S-13 | 2474 | 626.6 | 3619 | 6407 | 6407 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18333 | 40 | 8.692 | 200 | CMG-DM24 | A3081 | 2.864e-06 | 2 | 4.24e-07 | 25.6558 | 400.4215 | 0.63904 | 698324.0223 | -0.0016186 | -0.0022239 | 0.25 | 0.0035573 | 0.29052 | 0.0035304 | 0.29273 | 4 | 0.0035433 | 0.29074 | -0.36416 | -0.36416 | -0.68127 | -0.68127 | -3.7757 | 1.1137 | -3.4903 | 1.714
2017 | YKAR5 | .3077..HHZ | 2017-10-10 21:12 | 292.375 | 0.2 | 20 | 115 | -11.11 | -73 | 2e-05 | 0 | 20 | 0.27829 | 0 | 0.0036841 | S-13 | 2349 | 615.83 | 3762 | 6164 | 6164 | 0.1975 | 5 | 23 | 1 | 0.0205 | MAN-ELP-0110 | G18325 | 40 | 8.692 | 200 | CMG-DM24 | A3077 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 382.4276 | 0.621 | 697107.0059 | -0.0016895 | -0.0022948 | 0.25 | 0.0037312 | 0.27698 | 0.0036923 | 0.2799 | 4 | 0.0036841 | 0.27829 | 0.22115 | 0.22115 | -0.57517 | -0.57517 | -3.4423 | 1.0691 | -3.494 | 1.3257
2017 | YKAR6 | .3038..HHZ | 2017-10-10 21:14 | 292.475 | 0.2 | 20 | 116 | -10.81 | 107.3 | 1.6e-05 | 0 | 20 | 0.28807 | 0 | 0.0036017 | S-13 | 2350 | 616.95 | 3661 | 6211 | 6211 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18324 | 40 | 8.692 | 200 | CMG-DM24 | A3038 | 2.875e-06 | 2 | 4.24e-07 | 25.6558 | 388.156 | 0.62915 | 695652.1739 | -0.0016574 | -0.0022628 | 0.25 | 0.0036838 | 0.28054 | 0.0036501 | 0.28313 | 4 | 0.0036017 | 0.28807 | 1.3273 | 1.3273 | 1.7449 | 1.7449 | -5.8546 | -1.5079 | -3.6969 | 1.241
2017 | YKAR7 | .3074..HHZ | 2017-10-11 19:08 | 287.375 | 0.2 | 20 | 113 | -10.54 | -73.5 | 4.2e-05 | 2.5 | 20 | 0.29717 | 2.5 | 0.0035036 | S-13 | 2911 | 618.86 | 3577 | 6252 | 6252 | 0.1975 | 5 | 23 | 1 | 0.0248 | MAN-ELP-0110 | G18329 | 40 | 8.692 | 200 | CMG-DM24 | A3074 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 393.6426 | 0.63608 | 697836.7062 | -0.0016302 | -0.0022356 | 0.25 | 0.0036211 | 0.2854 | 0.003592 | 0.28771 | 4 | 0.0034841 | 0.29882 | 2.4616 | 3.0053 | 3.2858 | 3.8616 | -2.9677 | 2.0048 | -2.0414 | 3.4669
2017 | YKAR8 | .2165..HHZ | 2017-10-10 20:46 | 292.775 | 0.2 | 20 | 116 | -11.49 | -73 | 1.4e-05 | -0.2 | 20 | 0.26638 | -0.2 | 0.0038522 | S-13 | 1466 | 600.25 | 3720 | 5860 | 5860 | 0.1975 | 5 | 23 | 1 | 0.0211 | MAN-ELP-0110 | G18330 | 40 | 8.692 | 200 | CMG-DM24 | A2165 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.1675 | 0.61169 | 698080.2792 | -0.0017261 | -0.0023314 | 0.25 | 0.0038809 | 0.2663 | 0.0038347 | 0.2695 | 4 | 0.0038539 | 0.26626 | -0.45591 | -0.50071 | -1.1594 | -1.2035 | -1.6742 | 2.6426 | -2.3023 | 2.1195
2017 | YKAR9 | .3291..HHZ | 2017-10-10 20:48 | 292.975 | 0.2 | 20 | 116 | -11 | -72.6 | 1.7e-05 | -0.2 | 20 | 0.28184 | -0.2 | 0.0036697 | S-13 | 954 | 609.74 | 3688 | 6046 | 6046 | 0.1975 | 5 | 23 | 1 | 0.021 | MAN-ELP-0110 | G18322 | 40 | 8.692 | 200 | CMG-DM24 | A3291 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 378.7228 | 0.62112 | 697836.7062 | -0.001689 | -0.0022943 | 0.25 | 0.0037638 | 0.27458 | 0.0037246 | 0.27747 | 4 | 0.0036713 | 0.28171 | 1.473 | 1.4291 | 1.5728 | 1.5275 | -4.1535 | -0.030483 | -3.5695 | 1.0277
2018 | YKAB0 | .3292..HHZ | 2018-10-23 16:34 | 287.725 | 0.2 | 20 | 114 | -11.06 | -73.1 | 1.4e-05 | -1 | 20 | 0.2799 | -1 | 0.0036998 | S-13 | 2340 | 618.3 | 3754 | 6132 | 6132 | 0.1975 | 5 | 23 | 1 | 0.0075 | MAN-ELP-0110 | G18331 | 40 | 8.692 | 200 | CMG-DM24 | A3292 | 2.868e-06 | 2 | 4.24e-07 | 25.6558 | 383.5136 | 0.62027 | 697350.0697 | -0.0016923 | -0.0022977 | 0.25 | 0.0037194 | 0.27786 | 0.0036801 | 0.28083 | 4 | 0.0037081 | 0.27927 | -0.53518 | -0.75937 | -0.33054 | -0.5528 | -3.5542 | 0.20088 | -3.0286 | 0.96307
2018 | YKAB1 | .3085..HHZ | 2018-10-23 16:01 | 287.475 | 0.2 | 20 | 114 | -11.08 | 107.3 | 1.6e-05 | -1 | 20 | 0.27925 | -1 | 0.0036836 | S-13 | 2338 | 605.92 | 3729 | 5971 | 5971 | 0.1975 | 5 | 23 | 1 | 0.0212 | MAN-ELP-0110 | G18317 | 40 | 8.692 | 200 | CMG-DM24 | A3085 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 372.9844 | 0.61557 | 696864.1115 | -0.0017108 | -0.0023162 | 0.25 | 0.003827 | 0.27005 | 0.0037838 | 0.27313 | 4 | 0.0036918 | 0.27863 | 2.6484 | 2.4313 | 2.2424 | 2.0144 | -3.1344 | 0.80857 | -2.9855 | 1.4327
2018 | YKAB2 | .3087..HHZ | 2018-10-23 16:13 | 288.475 | 0.2 | 20 | 114 | -10.97 | -72.6 | 1.9e-05 | -1 | 20 | 0.28281 | -1 | 0.0036379 | S-13 | 2342 | 613.32 | 3831 | 6158 | 6158 | 0.1975 | 5 | 23 | 1 | 0.0275 | MAN-ELP-0110 | G18338 | 40 | 8.692 | 200 | CMG-DM24 | A3087 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 378.0984 | 0.61648 | 696864.1115 | -0.0017072 | -0.0023126 | 0.25 | 0.0037753 | 0.27375 | 0.0037332 | 0.27683 | 4 | 0.003646 | 0.28218 | 2.5519 | 2.3346 | 2.1597 | 1.9319 | -3.1314 | 0.81576 | -3.2255 | 1.1999
2018 | YKAB3 | .3080..HHZ | 2018-10-23 16:24 | 287.575 | 0.2 | 20 | 114 | -11.1 | -72.7 | 1.3e-05 | -1 | 20 | 0.27861 | -1 | 0.0036878 | S-13 | 2343 | 607.6 | 3764 | 5910 | 5910 | 0.1975 | 5 | 23 | 1 | 0.0055 | MAN-ELP-0110 | G18318 | 40 | 8.692 | 200 | CMG-DM24 | A3080 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 371.1925 | 0.61092 | 696864.1115 | -0.0017291 | -0.0023344 | 0.25 | 0.0038455 | 0.26875 | 0.0037993 | 0.27202 | 4 | 0.003696 | 0.27799 | 2.9345 | 2.718 | 2.4246 | 2.1962 | -0.99492 | 3.0305 | -1.9042 | 2.4739
2018 | YKAB4 | .3180..HHZ | 2018-10-23 16:26 | 287.975 | 0.2 | 20 | 114 | -10.8 | -72.5 | 1.3e-05 | -1 | 20 | 0.2884 | -1 | 0.0035938 | S-13 | 831 | 619.12 | 3646 | 6239 | 6239 | 0.1975 | 5 | 23 | 1 | 0.0195 | MAN-ELP-0110 | G18326 | 40 | 8.692 | 200 | CMG-DM24 | A3180 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 390.7627 | 0.63116 | 697107.0059 | -0.0016495 | -0.0022549 | 0.25 | 0.0036516 | 0.28302 | 0.0036194 | 0.28554 | 4 | 0.0036018 | 0.28776 | 0.7073 | 0.48588 | 1.004 | 0.77875 | -1.6522 | 2.2416 | -1.2873 | 2.7988
2018 | YKAB6 | .3036..HHZ | 2018-10-23 16:28 | 288.625 | 0.2 | 20 | 114 | -10.77 | 107.6 | 1.4e-05 | -1 | 20 | 0.2894 | -1 | 0.0035726 | S-13 | 2345 | 620.79 | 3746 | 6277 | 6277 | 0.1975 | 5 | 23 | 1 | 0.0226 | MAN-ELP-0110 | G18335 | 40 | 8.692 | 200 | CMG-DM24 | A3036 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 388.7757 | 0.62626 | 697107.0059 | -0.0016688 | -0.0022741 | 0.25 | 0.0036703 | 0.28158 | 0.003635 | 0.28431 | 4 | 0.0035806 | 0.28876 | 1.7175 | 1.4984 | 1.7912 | 1.5642 | -3.0701 | 0.76609 | -2.6111 | 1.6087
2018 | YKAB7 | .3086..HHZ | 2018-10-23 16:30 | 287.575 | 0.2 | 20 | 114 | -10.29 | -72.5 | 1.5e-05 | -1 | 20 | 0.30584 | -1 | 0.0033587 | S-13 | 2347 | 639.15 | 3957 | 6665 | 6665 | 0.1975 | 5 | 23 | 1 | 0.0242 | MAN-ELP-0110 | G18337 | 40 | 8.692 | 200 | CMG-DM24 | A3086 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 401.0483 | 0.62747 | 697107.0059 | -0.001664 | -0.0022694 | 0.25 | 0.003558 | 0.29047 | 0.0035245 | 0.29323 | 4 | 0.0033662 | 0.30516 | 4.7038 | 4.4913 | 4.3032 | 4.0706 | -2.5451 | 1.193 | -3.0457 | 1.294
2018 | YKAB8 | .4549..HHZ | 2018-10-23 16:31 | 286.975 | 0.2 | 20 | 113 | -10.9 | -72.6 | 1.5e-05 | -1 | 20 | 0.2851 | -1 | 0.0036001 | S-13 | 2493 | 621.46 | 3761 | 6296 | 6296 | 0.1975 | 5 | 23 | 1 | 0.0234 | MAN-ELP-0110 | G18336 | 40 | 8.692 | 200 | CMG-DM24 | A4549 | 2.873e-06 | 2 | 4.24e-07 | 25.6558 | 389.0536 | 0.62603 | 696136.4427 | -0.0016697 | -0.002275 | 0.25 | 0.0036728 | 0.28139 | 0.0036374 | 0.28413 | 4 | 0.0036081 | 0.28447 | 1.0247 | 0.804 | 0.34361 | 0.11984 | -3.2743 | 0.58113 | -3.5263 | 0.67407
2018 | YKAB9 | .3072..HHZ | 2018-10-23 16:33 | 288.525 | 0.2 | 20 | 114 | -11.15 | 107.3 | 1.4e-05 | -1 | 20 | 0.27701 | -1 | 0.003718 | S-13 | 2348 | 601.74 | 3754 | 5886 | 5886 | 0.1975 | 5 | 23 | 1 | 0.0207 | MAN-ELP-0110 | G18320 | 40 | 8.692 | 200 | CMG-DM24 | A3072 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.411 | 0.61058 | 698080.2792 | -0.0017304 | -0.0023358 | 0.25 | 0.0038783 | 0.26647 | 0.0038315 | 0.26973 | 4 | 0.0037263 | 0.2764 | 2.9624 | 2.746 | 2.7002 | 2.4712 | -2.7926 | 1.0051 | -2.6345 | 1.623
2018 | YKAR1 | .3079..HHZ | 2018-10-23 17:21 | 287.975 | 0.2 | 20 | 114 | -10.59 | -72.7 | 1.2e-05 | 0.1 | 20 | 0.29546 | 0.1 | 0.0034666 | S-13 | 1192 | 619.2 | 3774 | 6249 | 6249 | 0.1975 | 5 | 23 | 1 | 0.0233 | MAN-ELP-0110 | G18321 | 40 | 8.692 | 200 | CMG-DM24 | A3079 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0502 | 0.62347 | 697593.3031 | -0.0016798 | -0.0022851 | 0.25 | 0.0036936 | 0.2798 | 0.0036565 | 0.28264 | 4 | 0.0034658 | 0.29553 | 5.1939 | 5.215 | 4.5368 | 4.5602 | -1.3291 | 3.674 | -2.6816 | 3.0562
2018 | YKAR2 | .3076..HHZ | 2018-10-23 17:22 | 288.475 | 0.2 | 20 | 114 | -10.53 | -72.4 | 1.6e-05 | 0.1 | 20 | 0.29751 | 0.1 | 0.0034715 | S-13 | 832 | 628.26 | 3591 | 6355 | 6355 | 0.1975 | 5 | 23 | 1 | 0.0112 | MAN-ELP-0110 | G18332 | 40 | 8.692 | 200 | CMG-DM24 | A3076 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 401.4269 | 0.63895 | 697593.3031 | -0.0016189 | -0.0022243 | 0.25 | 0.0035522 | 0.29094 | 0.0035252 | 0.29316 | 4 | 0.0034707 | 0.29758 | 1.5244 | 1.5464 | 1.4824 | 1.505 | -2.7374 | 2.1309 | -2.4835 | 2.802
2018 | YKAR3 | .3073..HHZ | 2018-10-23 17:24 | 287.975 | 0.2 | 20 | 114 | -11.82 | -73 | 1.1e-05 | 0.1 | 20 | 0.25645 | 0.1 | 0.0040595 | S-13 | 822 | 619.92 | 3790 | 6255 | 6255 | 0.1975 | 5 | 23 | 1 | 0.0219 | MAN-ELP-0110 | G18334 | 40 | 8.692 | 200 | CMG-DM24 | A3073 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0229 | 0.6227 | 697593.3031 | -0.0016828 | -0.0022881 | 0.25 | 0.0036939 | 0.27978 | 0.0036563 | 0.28265 | 4 | 0.0040586 | 0.25651 | -11.027 | -11.0022 | -9.271 | -9.2508 | -6.3589 | -1.0357 | -2.0065 | 2.7081
2018 | YKAR4 | .3081..HHZ | 2018-10-23 17:25 | 287.475 | 0.2 | 20 | 114 | -10.66 | -72.8 | 1.3e-05 | 0.1 | 20 | 0.29309 | 0.1 | 0.0035265 | S-13 | 2474 | 626.6 | 3619 | 6407 | 6407 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18333 | 40 | 8.692 | 200 | CMG-DM24 | A3081 | 2.864e-06 | 2 | 4.24e-07 | 25.6558 | 400.4215 | 0.63904 | 698324.0223 | -0.0016186 | -0.0022239 | 0.25 | 0.0035573 | 0.29052 | 0.0035304 | 0.29273 | 4 | 0.0035257 | 0.29315 | 0.1117 | 0.13397 | 0.12238 | 0.14471 | -3.2999 | 1.6118 | -2.6866 | 2.54
2018 | YKAR5 | .3077..HHZ | 2018-10-23 17:27 | 287.525 | 0.2 | 20 | 114 | -11.02 | -72.8 | 1.9e-05 | 0.1 | 20 | 0.28119 | 0.1 | 0.0036814 | S-13 | 2349 | 615.83 | 3762 | 6164 | 6164 | 0.1975 | 5 | 23 | 1 | 0.0205 | MAN-ELP-0110 | G18325 | 40 | 8.692 | 200 | CMG-DM24 | A3077 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 382.4276 | 0.621 | 697107.0059 | -0.0016895 | -0.0022948 | 0.25 | 0.0037312 | 0.27698 | 0.0036923 | 0.2799 | 4 | 0.0036806 | 0.28125 | 0.29428 | 0.31651 | 0.46039 | 0.48279 | -3.3692 | 1.1645 | -2.4584 | 2.3836
2018 | YKAR6 | .3038..HHZ | 2018-10-23 17:27 | 287.4 | 0.2 | 20 | 113 | -10.7 | -72.5 | 1.5e-05 | 0.1 | 20 | 0.29174 | 0.1 | 0.0035438 | S-13 | 2350 | 616.95 | 3661 | 6211 | 6211 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18324 | 40 | 8.692 | 200 | CMG-DM24 | A3038 | 2.875e-06 | 2 | 4.24e-07 | 25.6558 | 388.156 | 0.62915 | 695652.1739 | -0.0016574 | -0.0022628 | 0.25 | 0.0036838 | 0.28054 | 0.0036501 | 0.28313 | 4 | 0.003543 | 0.29181 | 2.9136 | 2.9352 | 3.0416 | 3.0646 | -4.2683 | 0.1 | -2.4001 | 2.5607
2018 | YKAR7 | .3074..HHZ | 2018-10-23 17:28 | 287.825 | 0.2 | 20 | 114 | -10.69 | -73.2 | 1.1e-05 | 0.1 | 20 | 0.29208 | 0.1 | 0.0035308 | S-13 | 2911 | 618.86 | 3577 | 6252 | 6252 | 0.1975 | 5 | 23 | 1 | 0.0248 | MAN-ELP-0110 | G18329 | 40 | 8.692 | 200 | CMG-DM24 | A3074 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 393.6426 | 0.63608 | 697836.7062 | -0.0016302 | -0.0022356 | 0.25 | 0.0036211 | 0.2854 | 0.003592 | 0.28771 | 4 | 0.00353 | 0.29214 | 1.7043 | 1.7263 | 1.5174 | 1.5401 | -3.7249 | 0.72574 | -3.8098 | 1.1453
2018 | YKAR8 | .2165..HHZ | 2018-10-23 17:29 | 287.425 | 0.2 | 20 | 113 | -11.44 | 107 | 1.4e-05 | 0.1 | 20 | 0.26792 | 0.1 | 0.0038658 | S-13 | 1466 | 600.25 | 3720 | 5860 | 5860 | 0.1975 | 5 | 23 | 1 | 0.0211 | MAN-ELP-0110 | G18330 | 40 | 8.692 | 200 | CMG-DM24 | A2165 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.1675 | 0.61169 | 698080.2792 | -0.0017261 | -0.0023314 | 0.25 | 0.0038809 | 0.2663 | 0.0038347 | 0.2695 | 4 | 0.0038649 | 0.26798 | -0.81057 | -0.78808 | -0.58882 | -0.56665 | -2.0288 | 2.3553 | -1.7317 | 2.7564
2018 | YKAR9 | .3291..HHZ | 2018-10-23 17:30 | 286.975 | 0.2 | 20 | 113 | -10.91 | -72.5 | 1.7e-05 | 0.1 | 20 | 0.28477 | 0.1 | 0.0036074 | S-13 | 954 | 609.74 | 3688 | 6046 | 6046 | 0.1975 | 5 | 23 | 1 | 0.021 | MAN-ELP-0110 | G18322 | 40 | 8.692 | 200 | CMG-DM24 | A3291 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 378.7228 | 0.62112 | 697836.7062 | -0.001689 | -0.0022943 | 0.25 | 0.0037638 | 0.27458 | 0.0037246 | 0.27747 | 4 | 0.0036066 | 0.28484 | 3.1457 | 3.1673 | 2.6307 | 2.6536 | -2.4808 | 1.7077 | -2.5116 | 2.1538
2019 | YKAB0 | .3292..HHZ | 2019-03-06 16:02 | 287.975 | 0.2 | 20 | 114 | -10.81 | -72.6 | 1.2e-05 | -15.9 | 20 | 0.28807 | -15.9 | 0.0036024 | S-13 | 2340 | 618.3 | 3754 | 6132 | 6132 | 0.1975 | 5 | 23 | 1 | 0.0075 | MAN-ELP-0110 | G18331 | 40 | 8.692 | 200 | CMG-DM24 | A3292 | 2.868e-06 | 2 | 4.24e-07 | 25.6558 | 383.5136 | 0.62027 | 697350.0697 | -0.0016923 | -0.0022977 | 0.25 | 0.0037194 | 0.27786 | 0.0036801 | 0.28083 | 4 | 0.0037301 | 0.27786 | 2.1115 | -1.3593 | 2.5799 | -1.0573 | -0.90758 | -0.39909 | -0.11817 | 0.45856
2019 | YKAB1 | .3085..HHZ | 2019-03-06 16:21 | 280 | 0.2 | 20 | 111 | -10.79 | -72.2 | 1.2e-05 | -15.9 | 20 | 0.28874 | -15.9 | 0.0035837 | S-13 | 2338 | 605.92 | 3729 | 5971 | 5971 | 0.1975 | 5 | 23 | 1 | 0.0212 | MAN-ELP-0110 | G18317 | 40 | 8.692 | 200 | CMG-DM24 | A3085 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 372.9844 | 0.61557 | 696864.1115 | -0.0017108 | -0.0023162 | 0.25 | 0.003827 | 0.27005 | 0.0037838 | 0.27313 | 4 | 0.0037108 | 0.2785 | 5.2885 | 1.9304 | 5.7137 | 1.9654 | -0.49421 | 0.30767 | 0.48571 | 1.3837
2019 | YKAB2 | .3087..HHZ | 2019-03-06 16:40 | 288.975 | 0.2 | 20 | 114 | -10.68 | -72.1 | 2.1e-05 | -15.9 | 20 | 0.29242 | -15.9 | 0.0035247 | S-13 | 2342 | 613.32 | 3831 | 6158 | 6158 | 0.1975 | 5 | 23 | 1 | 0.0275 | MAN-ELP-0110 | G18338 | 40 | 8.692 | 200 | CMG-DM24 | A3087 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 378.0984 | 0.61648 | 696864.1115 | -0.0017072 | -0.0023126 | 0.25 | 0.0037753 | 0.27375 | 0.0037332 | 0.27683 | 4 | 0.0036497 | 0.28205 | 5.5842 | 2.2365 | 5.6281 | 1.8828 | -0.099112 | 0.71765 | 0.24294 | 1.1509
2019 | YKAB3 | .3080..HHZ | 2019-03-06 17:00 | 288.975 | 0.2 | 20 | 114 | -10.89 | -72.3 | 1.3e-05 | -14.3 | 20 | 0.28543 | -14.3 | 0.0036079 | S-13 | 2343 | 607.6 | 3764 | 5910 | 5910 | 0.1975 | 5 | 23 | 1 | 0.0055 | MAN-ELP-0110 | G18318 | 40 | 8.692 | 200 | CMG-DM24 | A3080 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 371.1925 | 0.61092 | 696864.1115 | -0.0017291 | -0.0023344 | 0.25 | 0.0038455 | 0.26875 | 0.0037993 | 0.27202 | 4 | 0.003723 | 0.27633 | 5.0375 | 2.0092 | 4.9311 | 1.585 | 1.1081 | 2.3217 | 0.6023 | 1.8627
2019 | YKAB4 | .3180..HHZ | 2019-03-06 17:20 | 287.975 | 0.2 | 20 | 114 | -10.63 | -72.2 | 1.2e-05 | -14.3 | 20 | 0.2941 | -14.3 | 0.0035114 | S-13 | 831 | 619.12 | 3646 | 6239 | 6239 | 0.1975 | 5 | 23 | 1 | 0.0195 | MAN-ELP-0110 | G18326 | 40 | 8.692 | 200 | CMG-DM24 | A3180 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 390.7627 | 0.63116 | 697107.0059 | -0.0016495 | -0.0022549 | 0.25 | 0.0036516 | 0.28302 | 0.0036194 | 0.28554 | 4 | 0.0036234 | 0.28472 | 2.9839 | -0.10982 | 3.0003 | -0.28427 | 0.62441 | 1.6459 | 0.70905 | 1.7357
2019 | YKAB6 | .3036..HHZ | 2019-03-06 17:40 | 287.975 | 0.2 | 20 | 114 | -10.55 | -72.1 | 1.6e-05 | -14.3 | 20 | 0.29682 | -14.3 | 0.0034708 | S-13 | 2345 | 620.79 | 3746 | 6277 | 6277 | 0.1975 | 5 | 23 | 1 | 0.0226 | MAN-ELP-0110 | G18335 | 40 | 8.692 | 200 | CMG-DM24 | A3036 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 388.7757 | 0.62626 | 697107.0059 | -0.0016688 | -0.0022741 | 0.25 | 0.0036703 | 0.28158 | 0.003635 | 0.28431 | 4 | 0.0035815 | 0.28736 | 4.5181 | 1.4732 | 4.4024 | 1.0731 | -0.2696 | 0.74096 | 0 | 1.1175
2019 | YKAB7 | .3086..HHZ | 2019-03-06 18:00 | 288.975 | 0.2 | 20 | 114 | -10.06 | 107.9 | 1.8e-05 | -12.3 | 20 | 0.31405 | -12.3 | 0.0032974 | S-13 | 2347 | 639.15 | 3957 | 6665 | 6665 | 0.1975 | 5 | 23 | 1 | 0.0242 | MAN-ELP-0110 | G18337 | 40 | 8.692 | 200 | CMG-DM24 | A3086 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 401.0483 | 0.62747 | 697107.0059 | -0.001664 | -0.0022694 | 0.25 | 0.003558 | 0.29047 | 0.0035245 | 0.29323 | 4 | 0.0033878 | 0.30544 | 6.4431 | 3.8769 | 7.102 | 4.1643 | -0.80579 | 0.57858 | -0.2469 | 1.3877
2019 | YKAB8 | .4549..HHZ | 2019-03-06 18:20 | 287.975 | 0.2 | 20 | 114 | -10.6 | -71.9 | 1.4e-05 | -12.3 | 20 | 0.29512 | -12.3 | 0.0034859 | S-13 | 2493 | 621.46 | 3761 | 6296 | 6296 | 0.1975 | 5 | 23 | 1 | 0.0234 | MAN-ELP-0110 | G18336 | 40 | 8.692 | 200 | CMG-DM24 | A4549 | 2.873e-06 | 2 | 4.24e-07 | 25.6558 | 389.0536 | 0.62603 | 696136.4427 | -0.0016697 | -0.002275 | 0.25 | 0.0036728 | 0.28139 | 0.0036374 | 0.28413 | 4 | 0.0035815 | 0.28703 | 4.1643 | 1.5357 | 3.8699 | 1.0208 | -0.13471 | 1.3128 | 0 | 1.5751
2019 | YKAB9 | .3072..HHZ | 2019-03-06 18:40 | 286.975 | 0.2 | 20 | 113 | -10.92 | 107.7 | 1.5e-05 | -12.3 | 20 | 0.28445 | -12.3 | 0.0036163 | S-13 | 2348 | 601.74 | 3754 | 5886 | 5886 | 0.1975 | 5 | 23 | 1 | 0.0207 | MAN-ELP-0110 | G18320 | 40 | 8.692 | 200 | CMG-DM24 | A3072 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.411 | 0.61058 | 698080.2792 | -0.0017304 | -0.0023358 | 0.25 | 0.0038783 | 0.26647 | 0.0038315 | 0.26973 | 4 | 0.0037155 | 0.27664 | 5.6167 | 3.0279 | 5.456 | 2.5635 | -0.13833 | 1.287 | 0.12134 | 1.7153
2019 | YKAR1 | .3079..HHZ | 2019-03-06 19:00 | 287.975 | 0.2 | 20 | 114 | -10.36 | 107.6 | 1.4e-05 | -10.8 | 20 | 0.30339 | -10.8 | 0.0033853 | S-13 | 1192 | 619.2 | 3774 | 6249 | 6249 | 0.1975 | 5 | 23 | 1 | 0.0233 | MAN-ELP-0110 | G18321 | 40 | 8.692 | 200 | CMG-DM24 | A3079 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0502 | 0.62347 | 697593.3031 | -0.0016798 | -0.0022851 | 0.25 | 0.0036936 | 0.2798 | 0.0036565 | 0.28264 | 4 | 0.0034668 | 0.29608 | 7.4173 | 5.1876 | 7.3419 | 4.7567 | 0.89429 | 3.6466 | 0.12351 | 3.2527
2019 | YKAR2 | .3076..HHZ | 2019-03-06 19:20 | 286.975 | 0.2 | 20 | 113 | -10.34 | -72.1 | 1.1e-05 | -10.8 | 20 | 0.30409 | -10.8 | 0.0033827 | S-13 | 832 | 628.26 | 3591 | 6355 | 6355 | 0.1975 | 5 | 23 | 1 | 0.0112 | MAN-ELP-0110 | G18332 | 40 | 8.692 | 200 | CMG-DM24 | A3076 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 401.4269 | 0.63895 | 697593.3031 | -0.0016189 | -0.0022243 | 0.25 | 0.0035522 | 0.29094 | 0.0035252 | 0.29316 | 4 | 0.0034642 | 0.29676 | 4.0434 | 1.7324 | 3.7267 | 1.2286 | -0.21842 | 2.3169 | -0.23911 | 2.5256
2019 | YKAR3 | .3073..HHZ | 2019-03-06 19:40 | 288.975 | 0.2 | 20 | 114 | -11.61 | -72.6 | 1.2e-05 | -10.8 | 20 | 0.26272 | -10.8 | 0.0039318 | S-13 | 822 | 619.92 | 3790 | 6255 | 6255 | 0.1975 | 5 | 23 | 1 | 0.0219 | MAN-ELP-0110 | G18334 | 40 | 8.692 | 200 | CMG-DM24 | A3073 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0229 | 0.6227 | 697593.3031 | -0.0016828 | -0.0022881 | 0.25 | 0.0036939 | 0.27978 | 0.0036563 | 0.28265 | 4 | 0.0040265 | 0.2564 | -7.5344 | -10.1242 | -7.0507 | -9.2893 | -2.8663 | -0.15772 | 0.21378 | 2.6696
2019 | YKAR4 | .3081..HHZ | 2019-03-06 20:02 | 287.975 | 0.2 | 20 | 114 | -10.46 | -72.5 | 1.6e-05 | -9.2 | 20 | 0.29992 | -9.2 | 0.0034397 | S-13 | 2474 | 626.6 | 3619 | 6407 | 6407 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18333 | 40 | 8.692 | 200 | CMG-DM24 | A3081 | 2.864e-06 | 2 | 4.24e-07 | 25.6558 | 400.4215 | 0.63904 | 698324.0223 | -0.0016186 | -0.0022239 | 0.25 | 0.0035573 | 0.29052 | 0.0035304 | 0.29273 | 4 | 0.0035103 | 0.29376 | 2.5703 | 0.57145 | 2.4545 | 0.35257 | -0.84125 | 2.0493 | -0.35448 | 2.7479
2019 | YKAR5 | .3077..HHZ | 2019-03-06 20:20 | 288.975 | 0.2 | 20 | 114 | -10.85 | 107.5 | 2.1e-05 | -9.2 | 20 | 0.28675 | -9.2 | 0.0036013 | S-13 | 2349 | 615.83 | 3762 | 6164 | 6164 | 0.1975 | 5 | 23 | 1 | 0.0205 | MAN-ELP-0110 | G18325 | 40 | 8.692 | 200 | CMG-DM24 | A3077 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 382.4276 | 0.621 | 697107.0059 | -0.0016895 | -0.0022948 | 0.25 | 0.0037312 | 0.27698 | 0.0036923 | 0.2799 | 4 | 0.0036752 | 0.28086 | 2.4637 | 0.46262 | 2.446 | 0.34418 | -1.1998 | 1.3106 | -0.47287 | 2.245
2019 | YKAR6 | .3038..HHZ | 2019-03-06 20:41 | 287.975 | 0.2 | 20 | 114 | -10.52 | -72.3 | 1.8e-05 | -9.2 | 20 | 0.29785 | -9.2 | 0.0034528 | S-13 | 2350 | 616.95 | 3661 | 6211 | 6211 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18324 | 40 | 8.692 | 200 | CMG-DM24 | A3038 | 2.875e-06 | 2 | 4.24e-07 | 25.6558 | 388.156 | 0.62915 | 695652.1739 | -0.0016574 | -0.0022628 | 0.25 | 0.0036838 | 0.28054 | 0.0036501 | 0.28313 | 4 | 0.0035236 | 0.29174 | 5.4066 | 3.4659 | 5.1992 | 3.0409 | -1.7753 | 0.63072 | -0.24251 | 2.5371
2019 | YKAR7 | .3074..HHZ | 2019-03-06 21:00 | 287.975 | 0.2 | 20 | 114 | -10.36 | -71.9 | 1.6e-05 | -8.4 | 20 | 0.30339 | -8.4 | 0.0033899 | S-13 | 2911 | 618.86 | 3577 | 6252 | 6252 | 0.1975 | 5 | 23 | 1 | 0.0248 | MAN-ELP-0110 | G18329 | 40 | 8.692 | 200 | CMG-DM24 | A3074 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 393.6426 | 0.63608 | 697836.7062 | -0.0016302 | -0.0022356 | 0.25 | 0.0036211 | 0.2854 | 0.003592 | 0.28771 | 4 | 0.0034534 | 0.29771 | 5.6269 | 3.8591 | 5.4486 | 3.4733 | 0.19766 | 2.8586 | 0.12133 | 3.0785
2019 | YKAR8 | .2165..HHZ | 2019-03-06 21:20 | 287.975 | 0.2 | 20 | 114 | -11.24 | -72.6 | 1.5e-05 | -8.4 | 20 | 0.27416 | -8.4 | 0.0037518 | S-13 | 1466 | 600.25 | 3720 | 5860 | 5860 | 0.1975 | 5 | 23 | 1 | 0.0211 | MAN-ELP-0110 | G18330 | 40 | 8.692 | 200 | CMG-DM24 | A2165 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.1675 | 0.61169 | 698080.2792 | -0.0017261 | -0.0023314 | 0.25 | 0.0038809 | 0.2663 | 0.0038347 | 0.2695 | 4 | 0.0038221 | 0.26902 | 2.1623 | 0.32958 | 1.7268 | -0.17878 | 0.94401 | 3.4729 | 0.5839 | 3.1442
2019 | YKAR9 | .3291..HHZ | 2019-03-06 21:40 | 286.975 | 0.2 | 20 | 113 | -10.75 | -72.2 | 2e-05 | -8.4 | 20 | 0.29007 | -8.4 | 0.003533 | S-13 | 954 | 609.74 | 3688 | 6046 | 6046 | 0.1975 | 5 | 23 | 1 | 0.021 | MAN-ELP-0110 | G18322 | 40 | 8.692 | 200 | CMG-DM24 | A3291 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 378.7228 | 0.62112 | 697836.7062 | -0.001689 | -0.0022943 | 0.25 | 0.0037638 | 0.27458 | 0.0037246 | 0.27747 | 4 | 0.0035992 | 0.28463 | 5.1433 | 3.3664 | 4.5388 | 2.5806 | -0.48328 | 1.9068 | -0.60351 | 2.0807
mean | YKAB0 | .3292..HHZ | 2015-11-23 21:00 | 274.925 | 0.2 | 20 | 108.6 | -10.98 | -73.04 | 1.14e-05 | -7.86 | 20 | 0.28253 | -7.86 | 0.003668 | S-13 | 2340 | 618.3 | 3754 | 6132 | 6132 | 0.1975 | 5 | 23 | 1 | 0.0075 | MAN-ELP-0110 | G18331 | 40 | 8.692 | 200 | CMG-DM24 | A3292 | 2.868e-06 | 2 | 4.24e-07 | 25.6558 | 383.5136 | 0.62027 | 697350.0697 | -0.0016923 | -0.0022977 | 0.25 | 0.0037194 | 0.27786 | 0.0036801 | 0.28083 | 4 | 0.0037311 | 0.2775 | 0.32893 | -1.3861 | 0.60567 | -1.186 | -2.6901 | -0.42584 | -2.0924 | 0.32988
mean | YKAB1 | .3085..HHZ | 2015-11-23 19:26 | 288.05 | 0.2 | 20 | 113.8 | -10.982 | 35.46 | 1.54e-05 | -8.22 | 20 | 0.28246 | -8.22 | 0.0036392 | S-13 | 2338 | 605.92 | 3729 | 5971 | 5971 | 0.1975 | 5 | 23 | 1 | 0.0212 | MAN-ELP-0110 | G18317 | 40 | 8.692 | 200 | CMG-DM24 | A3085 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 372.9844 | 0.61557 | 696864.1115 | -0.0017108 | -0.0023162 | 0.25 | 0.003827 | 0.27005 | 0.0037838 | 0.27313 | 4 | 0.0037049 | 0.2772 | 3.8223 | 2.0847 | 3.4168 | 1.4912 | -1.9605 | 0.46201 | -1.8112 | 0.90945
mean | YKAB2 | .3087..HHZ | 2015-11-23 19:36 | 290.385 | 0.2 | 20 | 114.8 | -10.868 | -36.44 | 1.96e-05 | -8.22 | 20 | 0.2862 | -8.22 | 0.0036074 | S-13 | 2342 | 613.32 | 3831 | 6158 | 6158 | 0.1975 | 5 | 23 | 1 | 0.0275 | MAN-ELP-0110 | G18338 | 40 | 8.692 | 200 | CMG-DM24 | A3087 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 378.0984 | 0.61648 | 696864.1115 | -0.0017072 | -0.0023126 | 0.25 | 0.0037753 | 0.27375 | 0.0037332 | 0.27683 | 4 | 0.0036723 | 0.28086 | 3.3679 | 1.63 | 3.3816 | 1.4554 | -2.3155 | 0.11109 | -2.0036 | 0.72338
mean | YKAB3 | .3080..HHZ | 2015-11-23 19:49 | 264.295 | 0.2 | 20 | 104.4 | -11.03 | -36.56 | 1.132e-05 | -7.82 | 20 | 0.28088 | -7.82 | 0.0036653 | S-13 | 2343 | 607.6 | 3764 | 5910 | 5910 | 0.1975 | 5 | 23 | 1 | 0.0055 | MAN-ELP-0110 | G18318 | 40 | 8.692 | 200 | CMG-DM24 | A3080 | 2.87e-06 | 2 | 4.24e-07 | 25.6558 | 371.1925 | 0.61092 | 696864.1115 | -0.0017291 | -0.0023344 | 0.25 | 0.0038455 | 0.26875 | 0.0037993 | 0.27202 | 4 | 0.0037288 | 0.27594 | 3.5256 | 1.8555 | 3.26 | 1.4405 | -0.40376 | 2.168 | -1.0688 | 1.7182
mean | YKAB4 | .3180..HHZ | 2015-11-23 23:30 | 289.185 | 0.2 | 20 | 114.4 | -10.748 | -36.42 | 1.3e-05 | -7.68 | 20 | 0.29015 | -7.68 | 0.0035574 | S-13 | 831 | 619.12 | 3646 | 6239 | 6239 | 0.1975 | 5 | 23 | 1 | 0.0195 | MAN-ELP-0110 | G18326 | 40 | 8.692 | 200 | CMG-DM24 | A3180 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 390.7627 | 0.63116 | 697107.0059 | -0.0016495 | -0.0022549 | 0.25 | 0.0036516 | 0.28302 | 0.0036194 | 0.28554 | 4 | 0.0036179 | 0.28514 | 1.7124 | 0.041096 | 1.6145 | -0.13934 | -0.64707 | 1.7968 | -0.6768 | 1.8807
mean | YKAB6 | .3036..HHZ | 2015-11-23 20:13 | 290.195 | 0.2 | 20 | 114.6 | -10.702 | -0.3 | 1.48e-05 | -7.68 | 20 | 0.29171 | -7.68 | 0.0035405 | S-13 | 2345 | 620.79 | 3746 | 6277 | 6277 | 0.1975 | 5 | 23 | 1 | 0.0226 | MAN-ELP-0110 | G18335 | 40 | 8.692 | 200 | CMG-DM24 | A3036 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 388.7757 | 0.62626 | 697107.0059 | -0.0016688 | -0.0022741 | 0.25 | 0.0036703 | 0.28158 | 0.003635 | 0.28431 | 4 | 0.0036001 | 0.28664 | 2.5995 | 0.95982 | 2.6026 | 0.82042 | -2.1882 | 0.22753 | -1.7997 | 0.86487
mean | YKAB7 | .3086..HHZ | 2015-11-23 20:25 | 289.295 | 0.2 | 20 | 114.2 | -10.206 | 35.64 | 1.6e-05 | -7.28 | 20 | 0.30885 | -7.28 | 0.0033445 | S-13 | 2347 | 639.15 | 3957 | 6665 | 6665 | 0.1975 | 5 | 23 | 1 | 0.0242 | MAN-ELP-0110 | G18337 | 40 | 8.692 | 200 | CMG-DM24 | A3086 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 401.0483 | 0.62747 | 697107.0059 | -0.001664 | -0.0022694 | 0.25 | 0.003558 | 0.29047 | 0.0035245 | 0.29323 | 4 | 0.003398 | 0.30376 | 5.1061 | 3.5894 | 5.3289 | 3.5939 | -2.1427 | 0.29105 | -2.02 | 0.81732
mean | YKAB8 | .3028..HHZ,.4549..HHZ | 2015-11-23 20:36 | 288.735 | 0.2 | 20 | 114.2 | -10.776 | -72.24 | 1.48e-05 | -7.28 | 20 | 0.28924 | -7.28 | 0.0035593 | S-13 | 2493 | 621.46 | 3761 | 6296 | 6296 | 0.1975 | 5 | 23 | 1 | 0.0234 | MAN-ELP-0110 | G18336 | 40 | 8.692 | 200 | CMG-DM24 | A4549 | 2.873e-06 | 2 | 4.24e-07 | 25.6558 | 389.0536 | 0.62603 | 696136.4427 | -0.0016697 | -0.002275 | 0.25 | 0.0036728 | 0.28139 | 0.0036374 | 0.28413 | 4 | 0.003616 | 0.28447 | 2.147 | 0.58622 | 1.801 | 0.12138 | -2.1521 | 0.36336 | -2.0689 | 0.67562
mean | YKAB9 | .3072..HHZ | 2015-11-23 20:47 | 288.755 | 0.2 | 20 | 113.8 | -11.066 | 71.44 | 1.42e-05 | -7.28 | 20 | 0.27973 | -7.28 | 0.0036801 | S-13 | 2348 | 601.74 | 3754 | 5886 | 5886 | 0.1975 | 5 | 23 | 1 | 0.0207 | MAN-ELP-0110 | G18320 | 40 | 8.692 | 200 | CMG-DM24 | A3072 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.411 | 0.61058 | 698080.2792 | -0.0017304 | -0.0023358 | 0.25 | 0.0038783 | 0.26647 | 0.0038315 | 0.26973 | 4 | 0.003739 | 0.27513 | 3.9511 | 2.4154 | 3.7079 | 2.0025 | -1.804 | 0.67446 | -1.6267 | 1.1543
mean | YKAR1 | .3079..HHZ | 2015-11-23 16:28 | 288.855 | 0.2 | 20 | 114.2 | -10.508 | -36.6 | 1.17e-05 | -7.86 | 20 | 0.29829 | -7.86 | 0.0034445 | S-13 | 1192 | 619.2 | 3774 | 6249 | 6249 | 0.1975 | 5 | 23 | 1 | 0.0233 | MAN-ELP-0110 | G18321 | 40 | 8.692 | 200 | CMG-DM24 | A3079 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0502 | 0.62347 | 697593.3031 | -0.0016798 | -0.0022851 | 0.25 | 0.0036936 | 0.2798 | 0.0036565 | 0.28264 | 4 | 0.0035044 | 0.29299 | 5.7978 | 4.161 | 5.5387 | 3.6636 | -0.72528 | 2.62 | -1.6797 | 2.1597
mean | YKAR2 | .3076..HHZ | 2015-11-23 17:05 | 290.015 | 0.2 | 20 | 114.6 | -10.462 | -72.26 | 1.278e-05 | -7.58 | 20 | 0.29987 | -7.58 | 0.0034373 | S-13 | 832 | 628.26 | 3591 | 6355 | 6355 | 0.1975 | 5 | 23 | 1 | 0.0112 | MAN-ELP-0110 | G18332 | 40 | 8.692 | 200 | CMG-DM24 | A3076 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 401.4269 | 0.63895 | 697593.3031 | -0.0016189 | -0.0022243 | 0.25 | 0.0035522 | 0.29094 | 0.0035252 | 0.29316 | 4 | 0.0034945 | 0.29474 | 2.4952 | 0.8716 | 2.2886 | 0.53682 | -1.7667 | 1.4561 | -1.6772 | 1.8338
mean | YKAR3 | .3073..HHZ | 2015-11-23 17:29 | 263.26 | 0.2 | 20 | 103.8 | -11.75 | -72.8 | 9.42e-06 | -7.58 | 20 | 0.25854 | -7.58 | 0.0039633 | S-13 | 822 | 619.92 | 3790 | 6255 | 6255 | 0.1975 | 5 | 23 | 1 | 0.0219 | MAN-ELP-0110 | G18334 | 40 | 8.692 | 200 | CMG-DM24 | A3073 | 2.867e-06 | 2 | 4.24e-07 | 25.6558 | 386.0229 | 0.6227 | 697593.3031 | -0.0016828 | -0.0022881 | 0.25 | 0.0036939 | 0.27978 | 0.0036563 | 0.28265 | 4 | 0.0040287 | 0.25412 | -8.3953 | -10.1848 | -8.5297 | -10.0945 | -3.7272 | -0.21835 | -1.2652 | 1.8643
mean | YKAR4 | .3081..HHZ | 2015-11-23 17:57 | 265.695 | 0.2 | 20 | 105 | -10.584 | -36.68 | 1.5e-05 | -7.26 | 20 | 0.29569 | -7.26 | 0.003487 | S-13 | 2474 | 626.6 | 3619 | 6407 | 6407 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18333 | 40 | 8.692 | 200 | CMG-DM24 | A3081 | 2.864e-06 | 2 | 4.24e-07 | 25.6558 | 400.4215 | 0.63904 | 698324.0223 | -0.0016186 | -0.0022239 | 0.25 | 0.0035573 | 0.29052 | 0.0035304 | 0.29273 | 4 | 0.0035426 | 0.29084 | 1.23 | -0.34308 | 1.0116 | -0.64663 | -2.1816 | 1.1348 | -1.7974 | 1.7487
mean | YKAR5 | .3077..HHZ | 2015-11-23 18:16 | 290.125 | 0.2 | 20 | 114.6 | -10.96 | 35.3 | 1.88e-05 | -6.92 | 20 | 0.28316 | -6.92 | 0.003635 | S-13 | 2349 | 615.83 | 3762 | 6164 | 6164 | 0.1975 | 5 | 23 | 1 | 0.0205 | MAN-ELP-0110 | G18325 | 40 | 8.692 | 200 | CMG-DM24 | A3077 | 2.869e-06 | 2 | 4.24e-07 | 25.6558 | 382.4276 | 0.621 | 697107.0059 | -0.0016895 | -0.0022948 | 0.25 | 0.0037312 | 0.27698 | 0.0036923 | 0.2799 | 4 | 0.0036902 | 0.27874 | 1.552 | 0.055535 | 1.1652 | -0.41645 | -2.1114 | 0.90349 | -1.7536 | 1.4844
mean | YKAR6 | .3038..HHZ | 2015-11-23 18:36 | 289.76 | 0.2 | 20 | 114.6 | -10.648 | -36.46 | 1.56e-05 | -6.92 | 20 | 0.29352 | -6.92 | 0.0034996 | S-13 | 2350 | 616.95 | 3661 | 6211 | 6211 | 0.1975 | 5 | 23 | 1 | 0.0244 | MAN-ELP-0110 | G18324 | 40 | 8.692 | 200 | CMG-DM24 | A3038 | 2.875e-06 | 2 | 4.24e-07 | 25.6558 | 388.156 | 0.62915 | 695652.1739 | -0.0016574 | -0.0022628 | 0.25 | 0.0036838 | 0.28054 | 0.0036501 | 0.28313 | 4 | 0.0035524 | 0.28893 | 4.125 | 2.6788 | 3.6701 | 2.0489 | -3.0569 | -0.15641 | -1.7716 | 1.545
mean | YKAR7 | .3074..HHZ | 2015-11-23 18:52 | 289.705 | 0.2 | 20 | 114.4 | -10.508 | -72.6 | 1.8e-05 | -6.24 | 20 | 0.2983 | -6.24 | 0.003457 | S-13 | 2911 | 618.86 | 3577 | 6252 | 6252 | 0.1975 | 5 | 23 | 1 | 0.0248 | MAN-ELP-0110 | G18329 | 40 | 8.692 | 200 | CMG-DM24 | A3074 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 393.6426 | 0.63608 | 697836.7062 | -0.0016302 | -0.0022356 | 0.25 | 0.0036211 | 0.2854 | 0.003592 | 0.28771 | 4 | 0.0035043 | 0.29408 | 3.7578 | 2.4424 | 3.6781 | 2.2145 | -1.6715 | 1.4419 | -1.6492 | 1.8197
mean | YKAR8 | .2165..HHZ | 2015-11-23 19:03 | 290.705 | 0.2 | 20 | 114.8 | -11.374 | -0.82 | 1.4e-05 | -6.54 | 20 | 0.26998 | -6.54 | 0.0038116 | S-13 | 1466 | 600.25 | 3720 | 5860 | 5860 | 0.1975 | 5 | 23 | 1 | 0.0211 | MAN-ELP-0110 | G18330 | 40 | 8.692 | 200 | CMG-DM24 | A2165 | 2.865e-06 | 2 | 4.24e-07 | 25.6558 | 367.1675 | 0.61169 | 698080.2792 | -0.0017261 | -0.0023314 | 0.25 | 0.0038809 | 0.2663 | 0.0038347 | 0.2695 | 4 | 0.0038668 | 0.266 | 0.60179 | -0.83602 | 0.1753 | -1.2991 | -0.61647 | 2.3073 | -0.96756 | 2.0239
mean | YKAR9 | .3291..HHZ | 2015-11-23 19:14 | 290.775 | 0.2 | 20 | 114.8 | -10.854 | -72.38 | 1.72e-05 | -6.54 | 20 | 0.28664 | -6.54 | 0.0035855 | S-13 | 954 | 609.74 | 3688 | 6046 | 6046 | 0.1975 | 5 | 23 | 1 | 0.021 | MAN-ELP-0110 | G18322 | 40 | 8.692 | 200 | CMG-DM24 | A3291 | 2.866e-06 | 2 | 4.24e-07 | 25.6558 | 378.7228 | 0.62112 | 697836.7062 | -0.001689 | -0.0022943 | 0.25 | 0.0037638 | 0.27458 | 0.0037246 | 0.27747 | 4 | 0.0036371 | 0.28241 | 3.7326 | 2.3496 | 3.3031 | 1.777 | -1.8939 | 0.89001 | -1.8392 | 1.2772

## reprocess_temperatures
| 2015 |  |  |  | 2016 |  |  |  | 2017 |  |  |  | 2018 |  |  |  | 2019 |  |  |  | Total G | Total A [%] | Total R [%] | Total T [°C]
| G | A [%] | R [%] | T [°C] | G | A [%] | R [%] | T [°C] | G | A [%] | R [%] | T [°C] | G | A [%] | R [%] | T [°C] | G | A [%] | R [%] | T [°C]
YKAB0 | 0.2884 | 2.698 | 0 | -18.4 | 0.27861 | -0.78848 | -0.11869 | -4 | 0.27765 | -1.1305 | 0.42801 | 0 | 0.2799 | -0.33054 | 0.99684 | -1 | 0.28807 | 2.5799 | 0.45988 | -15.9 | 0.28253 | 0.60567 | 0.35321 | -7.86
YKAB1 | 0.28741 | 5.2279 | 0 | -19.8 | 0.2799 | 2.4781 | 0.92154 | -4.4 | 0.27701 | 1.4217 | 0.88248 | 0 | 0.27925 | 2.2424 | 1.466 | -1 | 0.28874 | 5.7137 | 1.3782 | -15.9 | 0.28246 | 3.4168 | 0.92965 | -8.22
YKAB2 | 0.29174 | 5.3852 | 0 | -19.8 | 0.28347 | 2.3952 | 0.69182 | -4.4 | 0.28054 | 1.3397 | 0.65279 | 0 | 0.28281 | 2.1597 | 1.2358 | -1 | 0.29242 | 5.6281 | 1.1482 | -15.9 | 0.2862 | 3.3816 | 0.74572 | -8.22
YKAB3 | 0.28379 | 4.3288 | 0 | -19.8 | 0.27958 | 2.779 | 2.1572 | -4 | 0.27701 | 1.8367 | 2.1422 | 0 | 0.27861 | 2.4246 | 2.4951 | -1 | 0.28543 | 4.9311 | 1.856 | -14.3 | 0.28088 | 3.26 | 1.7301 | -7.82
YKAB4 | 0.29208 | 2.2913 | 0 | -18.9 | 0.28907 | 1.2368 | 2.3656 | -4 | 0.28708 | 0.53991 | 2.5346 | -0.2 | 0.2884 | 1.004 | 2.8131 | -1 | 0.2941 | 3.0003 | 1.7278 | -14.3 | 0.29015 | 1.6145 | 1.8882 | -7.68
YKAB6 | 0.29682 | 4.4024 | 0 | -19.1 | 0.2894 | 1.7912 | 0.95681 | -4 | 0.28609 | 0.62602 | 0.71258 | 0 | 0.2894 | 1.7912 | 1.6378 | -1 | 0.29682 | 4.4024 | 1.1175 | -14.3 | 0.29171 | 2.6026 | 0.88494 | -7.68
YKAB7 | 0.31477 | 7.3489 | 0 | -19.1 | 0.30655 | 4.5436 | 0.8658 | -4 | 0.30304 | 3.3469 | 0.61497 | 0 | 0.30584 | 4.3032 | 1.328 | -1 | 0.31405 | 7.102 | 1.3905 | -12.3 | 0.30885 | 5.3289 | 0.83984 | -7.28
YKAB8 | 0.29512 | 3.8699 | 0 | -19.1 | 0.28708 | 1.0392 | 0.72368 | -4 | 0.28379 | -0.11743 | 0.48126 | 0 | 0.2851 | 0.34361 | 0.71339 | -1 | 0.29512 | 3.8699 | 1.5751 | -12.3 | 0.28924 | 1.801 | 0.69868 | -7.28
YKAB9 | 0.28412 | 5.3347 | 0 | -19.1 | 0.27797 | 3.0556 | 1.3135 | -4 | 0.27511 | 1.9932 | 1.1823 | 0 | 0.27701 | 2.7002 | 1.6524 | -1 | 0.28445 | 5.456 | 1.714 | -12.3 | 0.27973 | 3.7079 | 1.1725 | -7.28
YKAR1 | 0.30304 | 7.2184 | 0 | -23.9 | 0.29648 | 4.8985 | 2.3677 | -4.5 | 0.29309 | 3.6978 | 2.1868 | -0.2 | 0.29546 | 4.5368 | 3.0861 | 0.1 | 0.30339 | 7.3419 | 3.2513 | -10.8 | 0.29829 | 5.5387 | 2.1784 | -7.86
YKAR2 | 0.30479 | 3.9658 | 0 | -22.7 | 0.29819 | 1.7163 | 2.0177 | -4.5 | 0.29478 | 0.55198 | 1.8871 | 0 | 0.29751 | 1.4824 | 2.8297 | 0.1 | 0.30409 | 3.7267 | 2.5283 | -10.8 | 0.29987 | 2.2886 | 1.8525 | -7.58
YKAR3 | 0.26212 | -7.2645 | 0 | -22.7 | 0.25674 | -9.1665 | 1.9021 | -4.5 | 0.25468 | -9.8956 | 2.0926 | 0 | 0.25645 | -9.271 | 2.7304 | 0.1 | 0.26272 | -7.0507 | 2.6672 | -10.8 | 0.25854 | -8.5297 | 1.8785 | -7.58
YKAR4 | 0.30095 | 2.809 | 0 | -22.7 | 0.29376 | 0.35319 | 1.7688 | -4.5 | 0.29074 | -0.68127 | 1.7529 | 0 | 0.29309 | 0.12238 | 2.57 | 0.1 | 0.29992 | 2.4545 | 2.7518 | -9.2 | 0.29569 | 1.0116 | 1.7687 | -7.26
YKAR5 | 0.28807 | 2.9188 | 0 | -21 | 0.28151 | 0.57611 | 1.4938 | -4.5 | 0.27829 | -0.57517 | 1.3646 | 0 | 0.28119 | 0.46039 | 2.4111 | 0.1 | 0.28675 | 2.446 | 2.2503 | -9.2 | 0.28316 | 1.1652 | 1.504 | -6.92
YKAR6 | 0.29854 | 5.4417 | 0 | -21 | 0.29141 | 2.923 | 1.4144 | -4.5 | 0.28807 | 1.7449 | 1.2822 | 0 | 0.29174 | 3.0416 | 2.5874 | 0.1 | 0.29785 | 5.1992 | 2.5398 | -9.2 | 0.29352 | 3.6701 | 1.5648 | -6.92
YKAR7 | 0.30304 | 5.3272 | 0 | -21 | 0.2958 | 2.8113 | 1.4358 | -4.4 | 0.29717 | 3.2858 | 3.4896 | 2.5 | 0.29208 | 1.5174 | 1.1878 | 0.1 | 0.30339 | 5.4486 | 3.0772 | -8.4 | 0.2983 | 3.6781 | 1.8381 | -6.24
YKAR8 | 0.27258 | 1.1429 | 0 | -19.8 | 0.26884 | -0.24487 | 2.1148 | -4.4 | 0.26638 | -1.1594 | 2.1452 | -0.2 | 0.26792 | -0.58882 | 2.7757 | 0.1 | 0.27416 | 1.7268 | 3.1377 | -8.4 | 0.26998 | 0.1753 | 2.0347 | -6.54
YKAR9 | 0.29174 | 5.1423 | 0 | -19.8 | 0.28477 | 2.6307 | 1.1519 | -4.4 | 0.28184 | 1.5728 | 1.0675 | -0.2 | 0.28477 | 2.6307 | 2.1818 | 0.1 | 0.29007 | 4.5388 | 2.0875 | -8.4 | 0.28664 | 3.3031 | 1.2977 | -6.54
Grand Total | 0.29217 | 3.7549 | 0 | -20.4278 | 0.28551 | 1.3904 | 1.4191 | -4.2778 | 0.28291 | 0.46656 | 1.4944 | 0.094444 | 0.28481 | 1.1428 | 2.0388 | -0.45 | 0.29231 | 3.8064 | 2.0366 | -11.8167 | 0.28754 | 2.1122 | 1.3978 | -7.3756

## manufacturer data
station_code | Kg | Rg | CDR | Rd | Sd | NCALIB0 | gain0
nominal | 629 | 3600 | 6300 | 6300 | 697837 | 0.003561 | 0.2902
YKAB0 | 618.3 | 3754 | 6132 | 6132 | 697350 | 0.003719 | 0.2779
YKAB1 | 605.92 | 3729 | 5971 | 5971 | 696864 | 0.003827 | 0.27
YKAB2 | 613.32 | 3831 | 6158 | 6158 | 696864 | 0.003775 | 0.2737
YKAB3 | 607.6 | 3764 | 5910 | 5910 | 696864 | 0.003846 | 0.2687
YKAB4 | 619.12 | 3646 | 6239 | 6239 | 697107 | 0.003652 | 0.283
YKAB6 | 620.79 | 3746 | 6277 | 6277 | 697107 | 0.00367 | 0.2816
YKAB7 | 639.15 | 3957 | 6665 | 6665 | 697107 | 0.003558 | 0.2905
YKAB8 | 621.46 | 3761 | 6296 | 6296 | 696136 | 0.003673 | 0.2814
YKAB9 | 601.74 | 3754 | 5886 | 5886 | 698080 | 0.003878 | 0.2665
YKAR1 | 619.2 | 3774 | 6249 | 6249 | 697593 | 0.003694 | 0.2798
YKAR2 | 628.26 | 3591 | 6355 | 6355 | 697593 | 0.003552 | 0.2909
YKAR3 | 619.92 | 3790 | 6255 | 6255 | 697593 | 0.003694 | 0.2798
YKAR4 | 626.6 | 3619 | 6407 | 6407 | 698324 | 0.003557 | 0.2905
YKAR5 | 615.83 | 3762 | 6164 | 6164 | 697107 | 0.003731 | 0.277
YKAR6 | 616.95 | 3661 | 6211 | 6211 | 695652 | 0.003684 | 0.2805
YKAR7 | 618.86 | 3577 | 6252 | 6252 | 697837 | 0.003621 | 0.2854
YKAR8 | 600.25 | 3720 | 5860 | 5860 | 698080 | 0.003881 | 0.2663
YKAR9 | 609.74 | 3688 | 6046 | 6046 | 697837 | 0.003764 | 0.2746
TEST1 | 621 | 3760 | 6225 | 6225 | 698324 | 0.003679 | 0.2809

## station IDs
station_code | sensor_id | preamp_id | digitizer_id
nominal | nominal | nominal | nominal
YKAB0 | 2340 | G18331 | A3292
YKAB1 | 2338 | G18317 | A3085
YKAB2 | 2342 | G18338 | A3087
YKAB3 | 2343 | G18318 | A3080
YKAB4 | 831 | G18326 | A3180
YKAB6 | 2345 | G18335 | A3036
YKAB7 | 2347 | G18337 | A3086
YKAB8 | 2493 | G18336 | A4549
YKAB9 | 2348 | G18320 | A3072
YKAR1 | 1192 | G18321 | A3079
YKAR2 | 832 | G18332 | A3076
YKAR3 | 822 | G18334 | A3073
YKAR4 | 2474 | G18333 | A3081
YKAR5 | 2349 | G18325 | A3077
YKAR6 | 2350 | G18324 | A3038
YKAR7 | 2911 | G18329 | A3074
YKAR8 | 1466 | G18330 | A2165
YKAR9 | 954 | G18322 | A3291
TEST1 | 2567 |  | A3082